## entry
Enter Project Specific Information here:
Company Owning Farm | DEFRA
JOB No. | AEAT/ED49056
Farm/Unit: | Site 5
Manufacturer: | Stack
Model (A/B if afterburner if fitted) | No Afterburner
Test 1 | Day 1
Test 2 | Day 2
Date of Site work Day 1 | 2001-07-24
Date of Site work Day 2 | 2001-07-25
Data Input By: | G Thistlethwaite
Date Inputted; | 2001-09-04
Moisture (Day1) | 6.5342
Moisture (Day2) | 8.2154
Enter Reference Conditions
Tref(oC)= | 0 | 273
Pref(KPa)= | 101.3
Oxygen ref, %= | 11 | (daily average) | Expression (ref o2): | 11
Moisture ref, %= | 0 | (daily average) | Expression (ref: H20): | 0
Oxygen duct, % = | 7.8774
Moisture duct, % = | No
Duct Diameter (m)= | 0.3 | Duct Shape = | round
Duct Area (m2)= | 0.070695
Internal Software Quality Assurance Plan
To be used for spreadsheets created for internal use only
Name: | Reporting Non-Particulate Results to EA
reference: | nonpart8.xlt
Date: | 2000-06-22
Purpose of Spreadsheet: | To calculate emission rate concentrations from stacks as part of contracts to the Environment Agency (EA). This should be used as an Appendix to reports for EA, the default setting is Appendix C, but this may be changed.
Responsible Officer: Alan Collings
Description:
Results areas: | Various results areas
Input areas: | In areas coloured yellow. Some coloured cells are already entered as they are generally appropriate but entries still need to be checked.
Other documentation: | EA Contract documents, hints page for help, EA Guidance Notes and plant authorisation notes.
Calculation sheets extracted from Calc_10.xlt
Checking:
Checking description: | All results areas checked.
Checking routines: | Initially all entries and calculations checked, followed by validation of changes only. The sheets are password protected for safety but may be unlocked by code holders to delete unwanted details only. Calculations may only be changed by the RO.
Name: Alan Collings Signature
Authorisation for use:
Name: Alan Collings Signature
Revisions:
Revision: | Issue 1
Date: | 1999-01-22
Checks: | Checked against Job number 20534020 by A Collings/N Watt
Checker Signature
Revision: | Issue 2
Changed C2 and cross references to spreadsheets
Date: | 1999-03-02
Checks: | Checked against Rechem Fawley (Feb 99) by A Collings
Checker Signature
Revision: | Issue 3
Changes | Introduced 26A sheet. Modified M29 sheet to include extra metals
Date: | 1999-05-13
Checks: | Checked against Cabot Carbon (March 1999) and EA Westbury (April 1999)
Checker Signature
Revision: | Issue 4
Changes | Changed layout of pitot data page and cross references from it.
Date: | 1999-07-16
Checks: | Checked with test data from SARP Killamash (20401020-June 1999)
Checker Signature
Revision: | Issue 5
Changes | Changed USEPA 26A sheets to adddress problems found in BS Tin Plate May 1999 report (AEAT 5641)
Date: | 1999-11-08
Checks: | Checked USEPA 26A sheets against BS Tin Plate May 1999 report (AEAT 5641) together with D Scott
Checker Signature
Revision: | Issue 6
Changes | Changes to refs.
Date: | 2000-01-16
Checks: | None
Checker Signature
Revision: | Issue 7
Changes | Cut down amounts of repeat entering and change in using reporting duct O2
Date: | 2000-03-01
Checks: | All changed spreadsheets checked using previous reports
Checker Signature
Revision: | Issue 8
Changes | Change calculation of VOCs expressed as toluene.
Date: | 2000-06-22
Checks: | All changed spreadsheets checked using previous reports
Checker Signature
Note on B21: Enter ref oxygen otherwise it will assume to be expressed at duct oxygen
Note on B23: This will take the average of the run 1 and run 2 oxygen (on-line spreadsheet) otherwise will assume 20.8999.

## summary
Customer | DEFRA |  | Date of Test: |  |  |  |  | 24-25/7/01
Site: | Site 5 |  | Plant identification/chimney: |  |  |  |  | Animal Incinerator
Manufacturer of Unit | Stack |  | Test carried out by: |  |  |  |  | B Donovan
Model Type (After burner fitted) | No Afterburner | [6 empty] | R Glass
Summary of Results
Velocity (actual), m/s = | 3.72 | (Day 1) |  |  | 4.43 | (Day 2)
Gas flowrate (actual), m3/s = | 0.26299 | (Day 1) |  |  | 0.31318 | (Day 2)
Duct pressure (velocity) | 99.6 | (Day 1) |  |  | 101.1 | (Day 2) |  |  | hide this row
Velocity (at ref. conditions *), m/s = | 1.6914 | (Day 1) |  |  | 1.8505 | (Day 2)
Gas flowrate (at ref. conditions *), m3/s = | 0.11957 | (Day 1) |  |  | 0.13082 | (Day 2)
Determinants | Expressed as | Concentration, C | [5 empty] | Mass emission rate, M
(kg/h)
|  | Day 1 |  |  | Day 2 |  |  | Day 1 |  | Day 2
Oxides of sulfur (absorption method): Run 1 | mg SO2 / m3 | 183.1604 | ± | 22.1583 | 262.984 | ± | 31.7409 |  | 0.078842 |  | 0.12385
Oxides of sulfur (absorption method): Run 2 | mg SO2 / m3 | 92.0795 | ± | 11.0814 | 30.1666 | ± | 3.6552 |  | 0.039636
Hydrogen chloride: Run 1 | mg HCl/ m3 | 13.5758 | ± | 1.6399 | 2.3195 | ± | 0.27977 |  | 0.0058438 |  | 0.0010924
Hydrogen chloride: Run 2 | mg HCl/ m3 | 40.5208 | ± | 4.87 | 47.8417 | ± | 5.7556 |  | 0.017442
Volatile organic compounds | mg C / m3 | 50.3594 | ± | 5.0359 | 105.1288 | ± | 10.5129 |  | 0.021677 |  | 0.04951
Carbon dioxide | % | 7.561 | ± | 0.7 | 7.4105 | ± | 0.7
Carbon monoxide | mg CO /m3 | 902.0372 | ± | 90.2037 | 1446.9394 | ± | 144.6939 |  | 0.38829 |  | 0.68143
Oxides of sulfur (by infra-red) | mg SO2 / m3 | 167.6851 | ± | 16.7685 | 84.6234 | ± | 8.4623 |  | 0.072181 |  | 0.039853
Oxides of nitrogen (by infra-red) | mg NO2 / m3 | 142.9071 | ± | 14.2907 | 113.9698 | ± | 11.397 |  | 0.061515 |  | 0.053674
Dioxins and Furans, total (minimum) | ng ITEQ/m3 | 0.32704 | ± | 0.078368 | 0.09038 | ± | 0.019073 |  | 0.00014078 |  | 4.2564e-05
Dioxins and Furans, total (maximum) | ng ITEQ/m3 | 0.32974 | ± | 0.078438 | 0.12341 | ± | 0.031842 |  | 0.00014194 |  | 5.8122e-05
Total particulate matter | mg/m3 | 99.8495 | ± | 9.985 | 245.2924 | ± | 24.5292 |  | 0.042981 |  | 0.11552
* Reference conditions | 101.3 | 0 | % H2O |  | 273 |  | 11 | % O2
Was a BS 3405 Particulate test carried out at the same time? |  | Yes
NOTES
1. All results arising from the monitoring of gaseous releases to atmosphere from a process are expressed at the reference conditions
specified for that process in the appropriate legislation as required by the regulatory authority.
2. Where the determined concentrations are less than the quoted measurement uncertainty, concentrations are reported as less than (<)
the measurement uncertainty.
3. Mass emission rates are calculated using the determined concentration or the quoted value for the measurement uncertainty,
whichever is the greater.

## methods
Project Title: | Assessment of Emissions and Environmental Impacts of Animal Incinerators in the UK
Customer Contract No: | DEFRA WA0806
AEAT Job Number | AEAT ED 49056
Site Name | Site 5
Methodology
Determinant | AEAT Procedure | Source Document | Other Details (Analysis. duration, no. of tests, equipment) | UKAS Accreditation
|  |  |  | Sampling | Analysis
Hydrogen chloride | AEAT/ENV/EQ/100 | EN 1911 (modified) | Absorption in deionised water. Analysis by Ion chromatography for chloride. | Yes | Yes
Oxides of sulfur | AEAT/ENV/EQ/100 | ISO 11632 | Absorption in 3% Hydrogen peroxide. Analysis by Ion chromatography for sulphate. | Yes | Yes
Dioxin and furans | AEAT/ENV/EQ/105 | EN 1948 (modified by single point sampling) | Multi-component sampling for 2,3,7,8 isomers for the length of a batch. Analysis by high resolution gas chromatography/high resolution mass spectrometry for toxic congeners. | Yes | Yes
Moisture | AEAT/ENV/EQ/115 | USEPA 4 (modified) | Absorption on silica gel and determination of absorbed water by gravimetry. | Yes | Yes
Carbon monoxide | AEAT/ENV/EQ/116 | ISO 12039 | Gas conditioning with on-line determination by infrared analyser for two test days. | Yes | Yes
Oxygen | AEAT/ENV/EQ/117 | ISO 12039 | Gas conditioning with on-line determination by paramagnetic analyser for two test days. | Yes | Yes
Total particulate matter (ISO 9096) | AEAT/ENV/EQ/118 | ISO 9096 | Isokinetic sampling with gravimetric determination. | Yes | Yes
Sulfur dioxide | AEAT/ENV/EQ/120 | ISO 7935 | Gas conditioning with on-line determination by infrared analyser for two test days. | Yes | Yes
Oxides of nitrogen | AEAT/ENV/EQ/121 | ISO 10849 | Gas conditioning with on-line determination by infrared analyser for two test days. | Yes | Yes
Carbon dioxide | AEAT/ENV/EQ/122 | ISO 12039 | Gas conditioning with on-line determination by infrared analyser for two test days. | Yes | Yes
Total organic carbon (VOCs) | AEAT/ENV/EQ/123 | EN 12619 | On-line determination by FID analyser for two test days. | Yes | Yes
Notes:
Checks were carried out to ensure that gas leaks were within the required tolerance of the standard methods used.
Equipment blanks were used to establish the level of contamination of hydrogen chloride, oxides of sulpur and dioxins and furans.
Sampling of on-line gases was carried out out a single point close to the mid point within the duct.
Sampling of total particulate matter and dioxins was carried out at four points within a single line across the duct.
Dust was removed prior to the analysis of CO, CO2, SO2, NOx and VOCs using a high efficency heated filter.
All analysers used have undergone tests to confirm that their performance characteristics meet the requirements of the standard methods
A sample of bottom ash was taken on each test day and submitted to UK Analytical Ltd for total organic carbon and total nitrogen analysis.
Samples of HCl, SO2 and dioxins and furans were analysed by Harwell Scientifics Ltd. (UKAS Lab Number 0322). Samples are archived at their analytical laboratory at
551 Harwell, Didcot, Oxon, OX11 0RA.
Deviations from Agreed Methods (Standard/AEAT)
A deviation from our UKAS accredited procedure was the sampling of particulates and dioxins across a single access line rather than two to allow simultaneous sampling
of particulates and dioxins. However, a velocity/temperature traverse indicated that the flow was uniform across the duct.

## Process
Project Title: | Assessment of Emissions and Environmental Impacts of Animal Incinerators in the UK
Customer Contract No: | DEFRA WA0806
AEAT Job Number | AEAT ED 49056
Site Name | Site 5
Process Data
Description of Process: | Breeding farm supplying farms with pigs.
Incineration of mainly pigs, but also sheep and game birds (depending on season)
Run burner for 2 hrs, blow for 2 hrs and repeat as required.
Date purchased: | 5 years ago
Frequency: | Batch |  | Length of operation: | up to 10h/day
Operation times? | Daily as required (approx 200 days/year), I.e 4 days per week
Feed stocks: | Production rate |  | Units
| Design Max. capacity | Actual
Pigs, sheep or birds | 350 | 250 | kg/batch
Operators assessment of annual throughput: |  | 250 kg/day four days a week
Provisions and condition for arrangement of feedstock storage: |  | No dedicated storage
Is plastic used or other material burned: | No
Fuels used: | Gas oil | (300 gallons per month)
Design/operating furnace temp. | 1000 o C
After burner fitted: | No | Design after burner temp. |  | n/a
Residence time (s) | 0.2 | Residence time (s) |  | n/a
Description Of Abatement systems: | None (although they have an irradiation burner)
Describe reagents used in abatement systems together with details of arrangements for their storage and handling: |  |  |  | None
Stack height (m): | Approx. 3 m from top of burner
Maintenance or Servicing | Twice per year or as required
Assessment of Condition of incinerator: | Poor
Standard of housekeeping on-site: | Good
Describe the site location: | Nearest houses are 1000m away (excluding farm houses)
Identify site sensitivities | Local houses and estate has two lakes
Residues
Bottom ash (removed 1-2 per week or as required)
Mixed with compost, allowed to settle for 3-6 months, then ploughed into the land
Noise and odour | Very little noise. Odour was present close by, but not noticeable above normal farm smell
within 20m.

## Moisture 1
Determination of Moisture |  |  |  | CLIENT | DEFRA
in Flue Gases by Absorption Method |  |  |  | JOB No. | AEAT/ED49056
| AEAT/ENV/EQ/WI | 115 |  | SITE: | Site 5
|  |  |  | DATE: | 2001-07-24
|  |  |  | Test: | Day 1
|  |  |  | Test period: | 1126-1328
|  |  |  | Meter Pressure, Pm (Kpa) |  |  | 100.7
|  |  |  | Meter Temperature, Tm (oC) |  |  | 23
Collector | Absorbent | Mass (g)
No. |  | Initial mass | Final mass | DW |  | Meter Readings (l)
1 | Drop Out | 260.1 | 271.8 | 11.7 |  | On | 1923.3
2 | Silica Gel | 456 | 460.5 | 4.5 |  | Off | 2182.6
3 | Silica Gel | 443.7 | 444.2 | 0.5 |  | DV | 259.3
Total mass of water collected, DW |  |  | = | 16.7
Volume Sampled (Vs)
Vs (STP)= | DV x | 273 | x | Pm | = | 237.7352
|  | 273+Tm |  | 101.3
A= | DW | x | 22.4 |  | = | 20.7822
18
% Moisture | = | 100 x A |  |  | = | 8.039
A + Vs
|  | % Moisture |  | = | 8.039 | ± | 0.8039
Data Input By: |  | G Thistlethwaite |  | Data Entry Checked By: |  |  | A Collings
Date: |  | 2001-09-04 |  | Date: |  |  | 2001-09-09

## Moisture 1 (2)
Determination of Moisture |  |  |  | CLIENT | DEFRA
in Flue Gases by Absorption Method |  |  |  | JOB No. | AEAT/ED49056
| AEAT/ENV/EQ/WI | 115 |  | SITE: | Site 5
|  |  |  | DATE: | 2001-07-24
|  |  |  | Test: | Day 1
|  |  |  | Test period: | 1332-1516
|  |  |  | Meter Pressure, Pm (Kpa) |  |  | 100.8
|  |  |  | Meter Temperature, Tm (oC) |  |  | 27
Collector | Absorbent | Mass (g)
No. |  | Initial mass | Final mass | DW |  | Meter Readings (l)
1 | Drop Out | 258.5 | 265.5 | 7 |  | On | 183.4
2 | Silica Gel | 449.2 | 452.1 | 2.9 |  | Off | 448.1
3 | Silica Gel | 455.9 | 456.2 | 0.3 |  | DV | 264.7
Total mass of water collected, DW |  |  | = | 10.2
Volume Sampled (Vs)
Vs (STP)= | DV x | 273 | x | Pm | = | 239.6881
|  | 273+Tm |  | 101.3
A= | DW | x | 22.4 |  | = | 12.6933
18
% Moisture | = | 100 x A |  |  | = | 5.0294
A + Vs
|  | % Moisture |  | = | 5.0294 | ± | 0.50294
Data Input By: |  | G Thistlethwaite |  | Data Entry Checked By: |  |  | A Collings
Date: |  | 2001-09-04 |  | Date: |  |  | 2001-09-09

## Moisture 2
Determination of Moisture |  |  |  | CLIENT | DEFRA
in Flue Gases by Absorption Method |  |  |  | JOB No. | AEAT/ED49056
| AEAT/ENV/EQ/WI | 115 |  | SITE: | Site 5
|  |  |  | DATE: | 2001-07-25
|  |  |  | Test: | Day 2
|  |  |  | Test period: | 1108-1311
|  |  |  | Meter Pressure, Pm (Kpa) |  |  | 100.8
|  |  |  | Meter Temperature, Tm (oC) |  |  | 30
Collector | Absorbent | Mass (g)
No. |  | Initial mass | Final mass | DW |  | Meter Readings (l)
1 | Drop Out | 271.8 | 281.9 | 10.1 |  | On | 449.1
2 | Silica Gel | 460.5 | 461.8 | 1.3 |  | Off | 629.8
3 | Silica Gel | 444.2 | 445.4 | 1.2 |  | DV | 180.7
Total mass of water collected, DW |  |  | = | 12.6
Volume Sampled (Vs)
Vs (STP)= | DV x | 273 | x | Pm | = | 162.0053
|  | 273+Tm |  | 101.3
A= | DW | x | 22.4 |  | = | 15.68
18
% Moisture | = | 100 x A |  |  | = | 8.8246
A + Vs
|  | % Moisture |  | = | 8.8246 | ± | 0.88246
Data Input By: |  | G Thistlethwaite |  | Data Entry Checked By: |  |  | A Collings
Date: |  | 2001-09-04 |  | Date: |  |  | 2001-09-09

## Moisture 2 (2)
Determination of Moisture |  |  |  | CLIENT | DEFRA
in Flue Gases by Absorption Method |  |  |  | JOB No. | AEAT/ED49056
| AEAT/ENV/EQ/WI | 115 |  | SITE: | Site 5
|  |  |  | DATE: | 2001-07-25
|  |  |  | Test: | Day 2
|  |  |  | Test period: | 1314-1500
|  |  |  | Meter Pressure, Pm (Kpa) |  |  | 100.8
|  |  |  | Meter Temperature, Tm (oC) |  |  | 31
Collector | Absorbent | Mass (g)
No. |  | Initial mass | Final mass | DW |  | Meter Readings (l)
1 | Drop Out | 265.5 | 271.6 | 6.1 |  | On | 630.7
2 | Silica Gel | 452.1 | 453.7 | 1.6 |  | Off | 772.8
3 | Silica Gel | 456.2 | 456.9 | 0.7 |  | DV | 142.1
Total mass of water collected, DW |  |  | = | 8.4
Volume Sampled (Vs)
Vs (STP)= | DV x | 273 | x | Pm | = | 126.9797
|  | 273+Tm |  | 101.3
A= | DW | x | 22.4 |  | = | 10.4533
18
% Moisture | = | 100 x A |  |  | = | 7.6061
A + Vs
|  | % Moisture |  | = | 7.6061 | ± | 0.76061
Data Input By: |  | G Thistlethwaite |  | Data Entry Checked By: |  |  | A Collings
Date: |  | 2001-09-04 |  | Date: |  |  | 2001-09-09

## On-line 1
Determination Of Concn. of |  | CLIENT | DEFRA |  |  |  |  | CLIENT | DEFRA
Combustion Gases Using On-Line |  | JOB No. | AEAT/ED49056 |  |  |  |  | JOB No. | AEAT/ED49056
Analysers |  | SITE: | Site 5 |  |  |  |  | SITE: | Site 5 |  |  |  |  | Standardised Concentrations
|  | TEST: | Day 1 |  |  |  |  | TEST: | Day 1 |  |  |  |  | Oxygen | 11
|  | DATE: | 2001-07-24 |  |  |  |  | DATE: | 2001-07-24 |  |  |  |  | Moisture | 0
| Determinant = | CO2 | CO | O2 | NOx | SO2 | VOCs | O2 | CO2 | CO | NOx | SO2 | VOC
Time | Units measured = | % | ppm | % | ppm | ppm | ppm | % | % | ppm | ppm | ppm | ppm
| % H2O of gas at analyser | 0 | 0 | 0 | 0 | 0 | 6.5342 | 0 | 0 | 0 | 0 | 0 | 6.5342
11:10 |  | 6.063 | 2290 | 12.821 | 119.14 | 3.126 | 133.52 | 12.821 | 7.4296 | 2806.1641 | 145.9941 | 3.8306 | 175.0537
11:15 |  | 6.3782 | 2310 | 12.4196 | 112.78 | 3.832 | 134.9 | 12.4196 | 7.4459 | 2696.6888 | 131.6591 | 4.4735 | 168.4916
11:20 |  | 6.8948 | 2270 | 11.7622 | 95.26 | 3.506 | 130.52 | 11.7622 | 7.4699 | 2459.3447 | 103.2058 | 3.7984 | 151.2927
11:25 |  | 6.973 | 2320 | 11.7004 | 87.74 | 3.226 | 127.82 | 11.7004 | 7.5039 | 2496.6303 | 94.42 | 3.4716 | 147.1677
11:30 |  | 6.844 | 2270 | 11.9316 | 105.04 | 3.072 | 100 | 11.9316 | 7.5549 | 2505.7981 | 115.9511 | 3.3911 | 118.1048
11:35 |  | 7.1228 | 2360 | 11.587 | 96.6 | 3.506 | 80.58 | 11.587 | 7.5718 | 2508.7512 | 102.6887 | 3.727 | 91.6474
11:40 |  | 7.2856 | 2280 | 11.3638 | 83.08 | 4.468 | 73.7 | 11.3638 | 7.5635 | 2366.9806 | 86.2494 | 4.6385 | 81.8606
11:45 |  | 7.9764 | 2540 | 10.5204 | 79.66 | 6.202 | 81.58 | 10.5204 | 7.6078 | 2422.6367 | 75.9792 | 5.9154 | 83.2503
11:50 |  | 9.1254 | 2260 | 9.036 | 72.2 | 4.644 | 72 | 9.036 | 7.6148 | 1885.8732 | 60.2478 | 3.8752 | 64.2812
11:55 |  | 9.6562 | 2100 | 8.262 | 68.46 | 5.468 | 67.74 | 8.262 | 7.5642 | 1645.0388 | 53.6283 | 4.2834 | 56.774
12:00 |  | 9.9386 | 1950 | 7.948 | 58.82 | 3.194 | 46.56 | 7.948 | 7.5967 | 1490.5034 | 44.9597 | 2.4414 | 38.0766
12:05 |  | 9.8854 | 1030 | 8.014 | 57.62 | 3.5 | 46.16 | 8.014 | 7.5947 | 791.3239 | 44.268 | 2.689 | 37.9429
12:10 |  | 9.7488 | 1000 | 8.182 | 54.16 | 4.112 | 37.54 | 8.182 | 7.5887 | 778.4243 | 42.1595 | 3.2009 | 31.265
12:15 |  | 9.4324 | 987.94 | 8.6172 | 54.3 | 14.482 | 31.68 | 8.6172 | 7.6026 | 796.2847 | 43.7661 | 11.6726 | 27.3193
12:20 |  | 9.375 | 721.82 | 8.7364 | 61.96 | 45.888 | 27.22 | 8.7364 | 7.6303 | 587.492 | 50.4295 | 37.3484 | 23.7033
12:25 |  | 9.2676 | 569 | 8.9238 | 74.14 | 79.856 | 24.76 | 8.9238 | 7.661 | 470.3579 | 61.2871 | 66.0121 | 21.8985
12:30 |  | 9.2612 | 534.54 | 8.9678 | 82.68 | 101.43 | 26.76 | 8.9678 | 7.6839 | 443.5013 | 68.5986 | 84.1552 | 23.7546
12:35 |  | 9.0778 | 564.92 | 9.1794 | 76.26 | 96.91 | 27.04 | 9.1794 | 7.6677 | 477.1691 | 64.4143 | 81.8566 | 24.4365
12:40 |  | 9.291 | 539.58 | 8.8608 | 72.78 | 85.046 | 25.38 | 8.8608 | 7.6401 | 443.7041 | 59.848 | 69.9345 | 22.3294
12:45 |  | 9.4326 | 553.74 | 8.688 | 77.06 | 97.624 | 25.84 | 8.688 | 7.6468 | 448.9048 | 62.4708 | 79.1416 | 22.4124
12:50 |  | 9.557 | 623.94 | 8.5304 | 76.6 | 107.856 | 22.78 | 8.5304 | 7.6489 | 499.3699 | 61.3068 | 86.3225 | 19.5066
12:55 |  | 9.655 | 657.18 | 8.416 | 76.42 | 110.944 | 22.64 | 8.416 | 7.6566 | 521.1536 | 60.6022 | 87.9803 | 19.209
13:00 |  | 9.8348 | 834.32 | 8.1792 | 75.94 | 115.55 | 22.56 | 8.1792 | 7.654 | 649.312 | 59.1005 | 89.9271 | 18.7848
13:05 |  | 10.2574 | 926.28 | 7.6344 | 70.58 | 126.544 | 24.04 | 7.6344 | 7.655 | 691.2746 | 52.6732 | 94.4387 | 19.1951
13:10 |  | 10.7554 | 904.94 | 6.9946 | 65.66 | 146.742 | 22.42 | 6.9946 | 7.6573 | 644.2753 | 46.7469 | 104.4735 | 17.0779
13:15 |  | 11.1966 | 1200 | 6.3804 | 61.7 | 137.852 | 24.1 | 6.3804 | 7.6343 | 818.2044 | 42.0693 | 93.9926 | 17.5811
13:20 |  | 10.8518 | 956.48 | 6.831 | 70.7 | 130.246 | 20.38 | 6.831 | 7.6361 | 673.0508 | 49.7498 | 91.6508 | 15.3435
13:25 |  | 10.8216 | 878.6 | 6.8912 | 78.64 | 131.6 | 17.4 | 6.8912 | 7.6476 | 620.9054 | 55.5748 | 93.0015 | 13.1562
13:30 |  | 10.8574 | 796.64 | 6.8456 | 83.04 | 134.818 | 15.16 | 6.8456 | 7.648 | 561.1578 | 58.4939 | 94.9666 | 11.4253
13:35 |  | 9.3396 | 869.88 | 8.8202 | 107.56 | 109.64 | 17.52 | 8.8202 | 7.6543 | 712.9101 | 88.1508 | 89.8555 | 15.3623
13:40 |  | 9.659 | 668.34 | 8.4412 | 107.14 | 93.46 | 15.06 | 8.4412 | 7.6752 | 531.0757 | 85.1355 | 74.2651 | 12.8036
13:45 |  | 10.101 | 644.12 | 7.8902 | 119.1 | 109.824 | 13.62 | 7.8902 | 7.6865 | 490.1527 | 90.6309 | 83.5722 | 11.0889
13:50 |  | 10.6298 | 652.1 | 7.238 | 107.88 | 127.134 | 12.98 | 7.238 | 7.7028 | 472.5362 | 78.1739 | 92.1261 | 10.0634
13:55 |  | 11.9154 | 833.76 | 5.587 | 82.18 | 164.556 | 12.18 | 5.587 | 7.7034 | 539.0338 | 53.1302 | 106.387 | 8.425
14:00 |  | 13.9024 | 2370 | 2.933 | 89.34 | 123.764 | 30.14 | 2.933 | 7.6604 | 1305.8941 | 49.2272 | 68.1952 | 17.7685
14:05 |  | 15.0014 | 2100 | 1.303 | 145.48 | 13.334 | 55.22 | 1.303 | 7.5784 | 1060.8767 | 73.4935 | 6.7361 | 29.8462
14:10 |  | 15.3052 | 1790 | 0.8204 | 172.56 | 13.264 | 54.76 | 0.8204 | 7.546 | 882.5375 | 85.0786 | 6.5397 | 28.8862
14:15 |  | 14.6716 | 1090 | 1.786 | 133.34 | 1.962 | 14.82 | 1.786 | 7.5991 | 564.56 | 69.0628 | 1.0162 | 8.2126
14:20 |  | 13.5724 | 763.68 | 3.266 | 98.84 | 13.896 | 9.68 | 3.266 | 7.6198 | 428.7417 | 55.4903 | 7.8014 | 5.8144
14:25 |  | 13.1384 | 446.98 | 3.8926 | 70.38 | 94.68 | 8.26 | 3.8926 | 7.6479 | 260.1869 | 40.9682 | 55.1132 | 5.1443
14:30 |  | 12.767 | 300.66 | 4.3388 | 64.74 | 123.892 | 6.92 | 4.3388 | 7.6319 | 179.7294 | 38.7005 | 74.0605 | 4.4259
14:35 |  | 11.9482 | 121.62 | 5.4134 | 64.64 | 116.8 | 6.1 | 5.4134 | 7.638 | 77.7471 | 41.3219 | 74.6658 | 4.1721
14:40 |  | 11.725 | 85.54 | 5.6082 | 62.82 | 130.95 | 5.46 | 5.6082 | 7.5908 | 55.3791 | 40.67 | 84.7778 | 3.782
14:45 |  | 11.5188 | 62.12 | 5.7398 | 64.5 | 156.514 | 7.94 | 5.7398 | 7.5221 | 40.566 | 42.1202 | 102.2077 | 5.5475
14:50 |  | 11.247 | 47.52 | 5.9896 | 69.44 | 154.41 | 10.7 | 5.9896 | 7.4676 | 31.5517 | 46.1058 | 102.523 | 7.6011
14:55 |  | 11.1676 | 43.54 | 5.9886 | 66.5 | 178.726 | 8.3 | 5.9886 | 7.4144 | 28.9071 | 44.1508 | 118.66 | 5.8958
15:00 |  | 10.885 | 29.94 | 6.2598 | 75.76 | 152.068 | 9.46 | 6.2598 | 7.3607 | 20.246 | 51.2304 | 102.8315 | 6.8443
15:05 |  | 10.3408 | 24.22 | 6.9828 | 89.72 | 123.472 | 10.58 | 6.9828 | 7.3559 | 17.2289 | 63.8223 | 87.8318 | 8.0522
15:10 |  | 10.1178 | 23.88 | 7.2376 | 94.38 | 152.724 | 12.46 | 7.2376 | 7.3315 | 17.3038 | 68.3893 | 110.6663 | 9.6599
15:15 |  | 8.8938 | 17.8 | 8.9204 | 82.84 | 135.604 | 8.56 | 8.9204 | 7.3499 | 14.71 | 68.4594 | 112.0638 | 7.5686
15:20 |  | 8.3818 | 14.42 | 9.6902 | 80.66 | 85.814 | 8.56 | 9.6902 | 7.4024 | 12.7351 | 71.2353 | 75.7871 | 8.0883
15:25 |  | 8.1792 | 12.6 | 9.9514 | 82.1 | 81.62 | 10.24 | 9.9514 | 7.3958 | 11.3932 | 74.2369 | 73.8029 | 9.9066
15:30 |  | 8.4216 | 12.18 | 9.6494 | 85.32 | 74.288 | 11.04 | 9.6494 | 7.4106 | 10.7178 | 75.0776 | 65.37 | 10.3938
15:35 |  | 8.2296 | 9.98 | 9.8594 | 85.16 | 78.484 | 11.72 | 9.8594 | 7.3794 | 8.949 | 76.3622 | 70.3758 | 11.2439
15:40 |  | 8.8398 | 8.96 | 9.2566 | 90.76 | 86.93 | 14.94 | 9.2566 | 7.5162 | 7.6184 | 77.1702 | 73.9137 | 13.5911
15:45 |  | 7.9832 | 6.74 | 10.3388 | 81.62 | 80.782 | 13.94 | 10.3388 | 7.4834 | 6.318 | 76.5101 | 75.7245 | 13.9808
15:50 |  | 6.6678 | 5.02 | 12.0574 | 75.4 | 72.484 | 11.08 | 12.0574 | 7.4651 | 5.6203 | 84.4163 | 81.1517 | 13.2722
15:55 |  | 6.6266 | 4.16 | 12.195 | 77.9 | 74.096 | 10.44 | 12.195 | 7.5363 | 4.7311 | 88.5939 | 84.2677 | 12.7032
16:00 |  | 7.2876 | 3.72 | 11.3004 | 93.88 | 41.258 | 10.68 | 11.3004 | 7.5157 | 3.8364 | 96.8178 | 42.5491 | 11.7842
16:05 |  | 6.7646 | 3.96 | 11.9452 | 105.46 | 29.286 | 10.48 | 11.9452 | 7.4786 | 4.378 | 116.5915 | 32.3772 | 12.3962
16:10 |  | 4.9532 | 7.32 | 14.294 | 90.68 | 17.516 | 9.2 | 14.294 | 7.4231 | 10.97 | 135.8965 | 26.2501 | 14.7514
Calibration
Determinant |  | CO2 | CO | O2 | NOx | SO2 | VOCs | O2 | CO2 | CO | NOx | SO2 | VOC
Average Concentration |  | 9.7213 | 857.4538 | 8.1839 | 84.7066 | 77.516 | 32.2266 | 8.1839 | 7.561 | 721.6298 | 69.5896 | 58.6898 | 31.3347
Max. Concentration |  | 15.3052 | 2540 | 14.294 | 172.56 | 178.726 | 134.9 | 14.294 | 7.7034 | 2806.1641 | 145.9941 | 118.66 | 175.0537
Min. Concentration |  | 4.9532 | 3.72 | 0.8204 | 54.16 | 1.962 | 5.46 | 0.8204 | 7.3315 | 3.8364 | 38.7005 | 1.0162 | 3.782
Units measured = |  | % | ppm | % | ppm | ppm | ppm propane | % | % | ppm | ppm | ppm | ppm propane
as mg/m3 at STP |  | n/a | 1071.8172 | n/a | 173.951 | 221.4743 | 51.7927 |  |  | 902.0372 | 142.9071 | 167.6851 | 50.3594
Expressed as |  | CO2 | as CO | as O2 | as NO2 | as SO2 | as C | as O2 | as CO2 | as CO | as NO2 | as SO2 | as C
at ref O2 (%) | 11 | n/a | 834.454 | n/a | 135.428 | 172.4269 | 40.3227
At ref H2O (%) | 0 | n/a | 834.454 | n/a | 135.428 | 172.4269 | 43.1417
Calibration Details
Time |  | 1015-1630 | [5 empty] | 1015-1630
Certified value |  | 14.8 | 750 | 19 | 151 | 392 | 76 | 19 | 14.8 | 750 | 151 | 392 | 76
Units |  | % | ppm | % | ppm | ppm | ppm | % | % | ppm | ppm | ppm | ppm
Zero Reading (start) |  | -0.02 | -0.1 | 0 | 0.3 | 0.1 | 0 | 0 | -0.02 | -0.1 | 0.3 | 0.1 | 0
Instrument Reading (start) |  | 14.8 | 755 | 19.2 | 150.4 | 392 | 76.1 | 19.2 | 14.8 | 755 | 150.4 | 392 | 76.1
Zero Reading (end) |  | -0.2 | 0.4 | 0.04 | 0.1 | -0.1 | 1.1 | 0.04 | -0.2 | 0.4 | 0.1 | -0.1 | 1.1
Instrument Reading (End) |  | 14.8 | 787 | 19.4 | 160.4 | 415 | 75 | 19.4 | 14.8 | 787 | 160.4 | 415 | 75
% Span Drift | % | 0 | 4.2384 | 1.0417 | 6.6489 | 5.8673 | -1.4455 | 1.0417 | 0 | 4.2384 | 6.6489 | 5.8673 | -1.4455
Acceptable Span Drift? |  | Pass | Fail | Pass | Fail | Fail | Pass | Pass | Pass | Pass | High | High | Pass
% Deviation from certified value | % | 1.2162 | 4.3067 | 1.2632 | 6.4901 | 5.8673 | 2.8947 | 1.2632 | 1.2162 | 4.3067 | 6.4901 | 5.8673 | 2.8947
Acceptable Deviation? |  | Pass | Pass | Pass | Fail | Fail | Pass | Pass | Pass | Pass | High | High | Pass
Substance used for calibration of FID for Totals VOCs= |  | Propane |  | Response Factor Used: |  | 3 |  | Propane |  | Response Factor Used: |  |  | 3
Make and Model of FID: |  | Bernath | [5 empty] | Bernath
Data Input By: | G Thistlethwaite |  |  | Data Entry Checked By: | [5 empty] | Data Entry Checked By: |  |  | A Collings
Date: | 2001-09-04 |  |  | Date: | [5 empty] | Date: |  |  | 2001-09-09

## On-line 2
Determination Of Concn. of |  | CLIENT | DEFRA |  |  |  |  | CLIENT | DEFRA
Combustion Gases Using On-Line |  | JOB No. | AEAT/ED49056 |  |  |  |  | JOB No. | AEAT/ED49056
Analysers |  | SITE: | Site 5 |  |  |  |  | SITE: | Site 5 |  |  |  |  | Standardised Concentrations
|  | TEST: | Day 2 |  |  |  |  | TEST: | Day 2 |  |  |  |  | Oxygen | 11
|  | DATE: | 2001-07-25 |  |  |  |  | DATE: | 2001-07-25 |  |  |  |  | Moisture | 0
| Determinant = | CO2 | CO | O2 | NOx | SO2 | VOCs | O2 | CO2 | CO | NOx | SO2 | VOC
Date/Time | Units measured = | ppm | ppm | % | ppm | ppm | ppm | % | % | ppm | ppm | ppm | ppm
| % H2O of gas at analyser | 0 | 0 | 0 | 0 | 0 | 8.2154 | 0 | 0 | 0 | 0 | 0 | 8.2154
10:45 |  | 7.3252 | 2330 | 11.2712 | 45 | 51.452 | 470.74 | 11.2712 | 7.5315 | 2395.6256 | 46.2674 | 52.9012 | 527.3199
10:50 |  | 9.5124 | 2260 | 6.9666 | 40.1 | 173.174 | 224.68 | 6.9666 | 6.7588 | 1605.7818 | 28.492 | 123.0441 | 173.9292
10:55 |  | 8.8288 | 1880 | 7.6052 | 36.4 | 202.828 | 220.44 | 7.6052 | 6.5744 | 1399.9458 | 27.1053 | 151.0363 | 178.8438
11:00 |  | 8.2042 | 1660 | 8.7836 | 32.62 | 142.978 | 137.08 | 8.7836 | 6.7034 | 1356.3435 | 26.653 | 116.8237 | 122.0297
11:05 |  | 8.518 | 1710 | 8.52 | 35.96 | 105.274 | 95.68 | 8.52 | 6.8116 | 1367.4475 | 28.7564 | 84.1852 | 83.3615
11:10 |  | 8.6816 | 1680 | 8.5382 | 44.82 | 69.31 | 140.42 | 8.5382 | 6.9527 | 1345.4351 | 35.8943 | 55.5072 | 122.5215
11:15 |  | 8.519 | 1700 | 9.0772 | 60.06 | 31.5 | 227 | 9.0772 | 7.1335 | 1423.5207 | 50.2921 | 26.377 | 207.0955
11:20 |  | 8.4142 | 1970 | 9.3796 | 66.26 | 18.532 | 155.74 | 9.3796 | 7.2307 | 1692.91 | 56.9402 | 15.9254 | 145.8135
11:25 |  | 8.3214 | 2105 | 9.569 | 66.66 | 15.024 | 114.08 | 9.569 | 7.2705 | 1839.1581 | 58.2415 | 13.1266 | 108.5942
11:30 |  | 8.2376 | 2220 | 9.8252 | 60.82 | 7.832 | 70.6 | 9.8252 | 7.3638 | 1984.5054 | 54.3683 | 7.0012 | 68.7597
11:35 |  | 8.0502 | 2360 | 10.1656 | 56.3 | 5.384 | 60.24 | 10.1656 | 7.4244 | 2176.5539 | 51.9237 | 4.9655 | 60.5302
11:40 |  | 8.3408 | 2550 | 9.771 | 53.2 | 5.634 | 61.78 | 9.771 | 7.4197 | 2268.3979 | 47.325 | 5.0118 | 59.8766
11:45 |  | 8.276 | 2480 | 9.8542 | 49.26 | 5.068 | 54.92 | 9.8542 | 7.4175 | 2222.7453 | 44.1502 | 4.5423 | 53.6289
11:50 |  | 8.2438 | 2550 | 9.9206 | 50.94 | 5.412 | 55.42 | 9.9206 | 7.4333 | 2299.306 | 45.932 | 4.8799 | 54.4444
11:55 |  | 8.5088 | 2450 | 9.5578 | 47.04 | 5.484 | 52.06 | 9.5578 | 7.4269 | 2138.474 | 41.0587 | 4.7867 | 49.5076
12:00 |  | 8.6722 | 2580 | 9.3342 | 46.24 | 5.554 | 50.46 | 9.3342 | 7.4232 | 2208.4075 | 39.5801 | 4.7541 | 47.0584
12:05 |  | 8.984 | 2440 | 8.8722 | 49.58 | 7.058 | 57.16 | 8.8722 | 7.3947 | 2008.3473 | 40.809 | 5.8094 | 51.2591
12:10 |  | 9.2912 | 2400 | 8.44 | 52.74 | 7.978 | 62.6 | 8.44 | 7.3823 | 1906.9021 | 41.9042 | 6.3389 | 54.1903
12:15 |  | 9.5136 | 2380 | 8.1254 | 56.8 | 8.35 | 58.7 | 8.1254 | 7.3728 | 1844.4413 | 44.0186 | 6.471 | 49.5628
12:20 |  | 9.5768 | 2280 | 8.1916 | 52.72 | 5.142 | 42.88 | 8.1916 | 7.4604 | 1776.1481 | 41.0695 | 4.0057 | 36.3939
12:25 |  | 10.0266 | 2390 | 7.68 | 58.82 | 4.408 | 41.88 | 7.68 | 7.5086 | 1789.7882 | 44.0483 | 3.301 | 34.1696
12:30 |  | 10.4274 | 2310 | 7.162 | 64.84 | 5.26 | 47.04 | 7.162 | 7.5143 | 1664.6528 | 46.7256 | 3.7905 | 36.9325
12:35 |  | 10.4096 | 2274 | 7.243 | 64.8 | 3.19 | 35.9 | 7.243 | 7.546 | 1648.4294 | 46.9737 | 2.3124 | 28.3533
12:40 |  | 10.577 | 2410 | 7.1238 | 61.48 | 2.266 | 30.6 | 7.1238 | 7.601 | 1731.8999 | 44.1814 | 1.6284 | 23.9584
12:45 |  | 11.0932 | 2320 | 6.4168 | 69.68 | 2.984 | 35.58 | 6.4168 | 7.5828 | 1585.8374 | 47.6298 | 2.0397 | 26.4976
12:50 |  | 11.4732 | 2500 | 5.801 | 90.42 | 4.122 | 34.78 | 5.801 | 7.5227 | 1639.1814 | 59.2859 | 2.7027 | 24.8454
12:55 |  | 11.334 | 2440 | 6.0614 | 92.44 | 1.946 | 26.2 | 6.0614 | 7.5618 | 1627.9164 | 61.674 | 1.2983 | 19.0447
13:00 |  | 12.051 | 2400 | 5.1124 | 114.52 | 5.132 | 47.24 | 5.1124 | 7.5569 | 1504.9786 | 71.8126 | 3.2181 | 32.2745
13:05 |  | 12.3864 | 1860 | 4.6434 | 140.9 | 4.468 | 37.86 | 4.6434 | 7.5431 | 1132.7092 | 85.8058 | 2.7209 | 25.1198
13:10 |  | 12.1192 | 1600 | 5.089 | 118.78 | 0.738 | 21.96 | 5.089 | 7.5884 | 1001.8342 | 74.3737 | 0.4621 | 14.9809
13:15 |  | 11.6326 | 962.46 | 5.656 | 103.7 | -0.096 | 17.12 | 5.656 | 7.5546 | 625.056 | 67.3465 | -0.062346 | 12.1135
13:20 |  | 8.254 | 302.24 | 10.6448 | 55.88 | -1.028 | 12.44 | 10.6448 | 7.9681 | 291.7716 | 53.9445 | -0.99239 | 13.084
13:25 |  | 11.1538 | 639.42 | 6.4776 | 61.4 | -0.96 | 12.76 | 6.4776 | 7.6563 | 438.9185 | 42.1469 | -0.65897 | 9.5429
13:30 |  | 11.3462 | 714.6 | 6.1822 | 75.8 | -1.082 | 11.86 | 6.1822 | 7.6321 | 480.6792 | 50.9872 | -0.72781 | 8.6917
13:35 |  | 11.2654 | 714.06 | 6.237 | 81.86 | 0.392 | 10.84 | 6.237 | 7.606 | 482.111 | 55.2693 | 0.26467 | 7.9739
13:40 |  | 10.993 | 611.58 | 6.527 | 81.52 | 11.574 | 10.36 | 6.527 | 7.5719 | 421.2511 | 56.1503 | 7.9721 | 7.7746
13:45 |  | 10.6906 | 533.74 | 6.9054 | 80.74 | 36.192 | 9.86 | 6.9054 | 7.5627 | 377.5761 | 57.1167 | 25.6028 | 7.5994
13:50 |  | 10.4154 | 476.84 | 7.2192 | 78.94 | 66.016 | 9.38 | 7.2192 | 7.537 | 345.0614 | 57.1243 | 47.7719 | 7.3953
13:55 |  | 10.1928 | 391.72 | 7.487 | 76.52 | 85.242 | 8.92 | 7.487 | 7.5232 | 289.1246 | 56.4786 | 62.9163 | 7.1731
14:00 |  | 10.1744 | 349.66 | 7.4996 | 81.24 | 99.128 | 9.04 | 7.4996 | 7.5167 | 258.3232 | 60.0188 | 73.2342 | 7.2764
14:05 |  | 10.2786 | 465.96 | 7.34 | 102.42 | 142.1 | 9.3 | 7.34 | 7.5043 | 340.192 | 74.7757 | 103.7456 | 7.3976
14:10 |  | 10.5118 | 866.14 | 6.9866 | 157.44 | 185.086 | 15.12 | 6.9866 | 7.4796 | 616.297 | 112.0255 | 131.6969 | 11.7215
14:15 |  | 11.1036 | 2600 | 6.0466 | 264.94 | 225.982 | 75 | 6.0466 | 7.4007 | 1732.9366 | 176.5862 | 150.6202 | 54.4629
14:20 |  | 11.4436 | 2450 | 5.4684 | 175.36 | 173.878 | 264.32 | 5.4684 | 7.3415 | 1571.7748 | 112.5006 | 111.5498 | 184.7499
14:25 |  | 11.6214 | 2209 | 5.1916 | 102.36 | 73.374 | 353.02 | 5.1916 | 7.3242 | 1392.1914 | 64.511 | 46.2429 | 242.4
14:30 |  | 11.8466 | 2080 | 4.8494 | 54.92 | 74.826 | 412.76 | 4.8494 | 7.307 | 1282.9427 | 33.8746 | 46.1526 | 277.3777
14:35 |  | 11.3746 | 1780 | 5.4192 | 70.96 | 49.78 | 259.64 | 5.4192 | 7.2741 | 1138.3133 | 45.3791 | 31.8344 | 180.902
14:40 |  | 11.1168 | 1560 | 5.8564 | 98.44 | 29.868 | 140.64 | 5.8564 | 7.3158 | 1026.616 | 64.7821 | 19.6557 | 100.8375
14:45 |  | 10.5718 | 1100 | 6.6454 | 133.96 | 4.388 | 31.3 | 6.6454 | 7.3422 | 763.9639 | 93.0369 | 3.0475 | 23.684
14:50 |  | 10.2234 | 556.36 | 7.3086 | 110.08 | -1.208 | 12.78 | 7.3086 | 7.4467 | 405.2536 | 80.1825 | 0 | 10.1422
14:55 |  | 10.0084 | 199.62 | 7.6248 | 75.94 | 19.934 | 10.28 | 7.6248 | 7.4638 | 148.8669 | 56.6324 | 14.8658 | 8.3525
15:00 |  | 10.3326 | 67.2 | 7.2512 | 58.26 | 69.186 | 10.84 | 7.2512 | 7.4946 | 48.7427 | 42.2582 | 50.1833 | 8.5664
15:05 |  | 10.9456 | 53.56 | 6.4406 | 63.24 | 60.234 | 12.82 | 6.4406 | 7.4942 | 36.6712 | 43.2989 | 41.2408 | 9.5632
15:10 |  | 10.376 | 24.18 | 7.1292 | 55.72 | 19.364 | 13.58 | 7.1292 | 7.4594 | 17.3833 | 40.0578 | 13.921 | 10.6367
15:15 |  | 9.6866 | 10.52 | 8.0636 | 56.46 | 6.718 | 13.94 | 8.0636 | 7.4707 | 8.1135 | 43.5445 | 5.1812 | 11.7134
15:20 |  | 9.4046 | 7.62 | 8.4726 | 57.36 | 6.784 | 13.6 | 8.4726 | 7.492 | 6.0703 | 45.6945 | 5.4043 | 11.8039
15:25 |  | 9.11 | 4.22 | 8.9112 | 57.66 | 6.688 | 17.26 | 8.9112 | 7.5228 | 3.4848 | 47.6139 | 5.5228 | 15.5285
15:30 |  | 8.9284 | 0.8 | 9.1696 | 61.94 | 0.832 | 15.9 | 9.1696 | 7.5352 | 0.67517 | 52.2749 | 0.70218 | 14.6201
Determinant |  | CO2 | CO | O2 | NOx | SO2 | VOCs | O2 | CO2 | CO | NOx | SO2 | VOC
Average Concentration |  | 9.9814 | 1521.0431 | 7.5709 | 76.1252 | 40.6323 | 78.731 | 7.5709 | 7.4105 | 1157.5515 | 55.4984 | 29.6182 | 65.4135
Max. Concentration |  | 12.3864 | 2600 | 11.2712 | 264.94 | 225.982 | 470.74 | 11.2712 | 7.9681 | 2395.6256 | 176.5862 | 151.0363 | 527.3199
Min. Concentration |  | 7.3252 | 0.8 | 4.6434 | 32.62 | -1.208 | 8.92 | 4.6434 | 6.5744 | 0.67517 | 26.653 | -0.99239 | 7.1731
Units measured = |  | % | ppm | % | ppm | ppm | ppm propane | % | % | ppm | ppm | ppm | ppm propane
as mg/m3 at STP |  | n/a | 1901.3039 | n/a | 156.3285 | 116.0923 | 126.532 |  |  | 1446.9394 | 113.9698 | 84.6234 | 105.1288
Expressed as |  | CO2 | as CO | as O2 | as NO2 | as SO2 | as C | as O2 | as CO2 | as CO | as NO2 | as SO2 | as C
at ref O2 (%) | 11 | n/a | 1412.1661 | n/a | 116.1107 | 86.2259 | 93.9798
At ref H2O (%) | 0 | n/a | 1412.1661 | n/a | 116.1107 | 86.2259 | 102.3917
Calibration Details
Time |  | 1040-1624 | [5 empty] | 1040-1624
Certified value |  | 14.8 | 750 | 19 | 151 | 392 | 76 | 19 | 14.8 | 750 | 151 | 392 | 76
Units |  | % | ppm | % | ppm | ppm | ppm | % | % | ppm | ppm | ppm | ppm
Zero Reading (start) |  | -0.2 | 0.2 | 0 | 0.2 | 0.1 | 0 | 0 | -0.2 | 0.2 | 0.2 | 0.1 | 0
Instrument Reading (start) |  | 14.8 | 756.4 | 19.2 | 150.8 | 390.7 | 76.3 | 19.2 | 14.8 | 756.4 | 150.8 | 390.7 | 76.3
Zero Reading (end) |  | -0.1 | 2.9 | 0 | 4.9 | -2.3 | 0.3 | 0 | -0.1 | 2.9 | 4.9 | -2.3 | 0.3
Instrument Reading (End) |  | 14.8 | 778 | 19.3 | 153.8 | 389.1 | 75 | 19.3 | 14.8 | 778 | 153.8 | 389.1 | 75
% Span Drift | % | 0 | 2.8556 | 0.52083 | 1.9894 | -0.40952 | -1.7038 | 0.52083 | 0 | 2.8556 | 1.9894 | -0.40952 | -1.7038
Acceptable Span Drift? |  | Pass | Fail | Pass | Pass | Pass | Pass | Pass | Pass | Pass | Pass | Pass | Pass
% Deviation from certified value | % | -0.67568 | 3.24 | 0.52632 | 5.0993 | 0.96939 | 2.1053 | 0.52632 | -0.67568 | 3.24 | 5.0993 | 0.96939 | 2.1053
Acceptable Deviation? |  | Pass | Pass | Pass | Fail | Pass | Pass | Pass | Pass | Pass | Fail | Pass | Pass
Substance used for calibration of FID for Totals VOCs= |  | Propane |  | Response Factor Used: |  | 3 |  | Propane |  | Response Factor Used: |  |  | 3
Make and Model of FID: |  | Bernath | [5 empty] | Bernath
Data Input By: | G Thistlethwaite |  |  | Data Entry Checked By: | [5 empty] | Data Entry Checked By: |  |  | A Collings
Date: | 2001-09-04 |  |  | Date: | [5 empty] | Date: |  |  | 2001-09-09

## SOx 1
Determination of Pollutants |  |  |  | CLIENT | DEFRA |  |  |  | Duct Oxygen, O2m | 8.8952 | %
in Flue Gases by Absorption Method |  |  |  | JOB No. | AEAT/ED49056
| AEAT/ENV/EQ/WI | 100 |  | SITE: | Site 5 |  |  |  |  | Vm | Tm | Pm | O2m | Samp Vol | Samp Con | Initial Vol | Blank Con
Pollutant | SO2 | Mw(p) = | 64 | TEST: | Day 1 |  |  |  | Value | 242.4 | 296 | 100.7 | 8.8952 | 309 | 240 | 200 | 0.6
Analyte | SO4 | Mw(a) = | 96 | DATE: | 2001-07-24 |  |  |  | Abs Error | 4.848 | 2.5 | 1 | 0.2 | 3.09 | 24 | 2 | 0.4
Ratio analyte/pollutant |  | K = | 1 | TEST PERIOD: | 1123-1318 |  |  |  | u | 2.799 | 1.4434 | 0.57735 | 0.11547 | 1.784 | 13.8564 | 1.1547 | 0.23094
|  | Ref Temperature, Tr | 0 | oC K | 273
|  | Ref Pressure, Pr | 101.3 | kPa |  |  |  | Mw(p) = | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64
|  | Ref Oxygen, O2r | 11 | % |  | Uncertainty |  | Mw(a) = | 96 | 96 | 96 | 96 | 96 | 96 | 96 | 96 | 96
|  | Ref Moisture, H2Or | 0 | % |  | % | Absolute | K = | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
|  | Meter Reading (start) | 160.6 | dm3
|  | Meter Reading (end) | 403 | dm3
|  | Metered Gas Volume, Vm | 242.4 | dm3 |  | 2 | 4.848 | Tr | 273 | 273 | 273 | 273 | 273 | 273 | 273 | 273 | 273
|  | Meter Temperature, Tm | 23 | oC K | 296 |  | 2.5 | Pr | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3
|  | Meter Pressure, Pm | 100.7 | kPa |  |  | 1 | O2r | 11 | 11 | 11 | 11 | 11 | 11 | 11 | 11 | 11
|  | Duct Oxygen, O2duct | 8.8952 | % |  |  | 0.2 | H2Or | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | Meter Moisture, H2Om | 0 | % |  |  |  | Vm | 242.4 | 245.199 | 242.4 | 242.4 | 242.4 | 242.4 | 242.4 | 242.4 | 242.4
|  | LOD | 0.2 | ug/ml |  |  |  | Tm | 296 | 296 | 297.4434 | 296 | 296 | 296 | 296 | 296 | 296
Sample id |  | Sample Reagent Volume | 309 | ml |  | 1 | 3.09 | Pm | 100.7 | 100.7 | 100.7 | 101.2774 | 100.7 | 100.7 | 100.7 | 100.7 | 100.7
PE/60448 |  | Sample Conc | 240 | ug/ml |  | 10 | 24 | O2m | 8.8952 | 8.8952 | 8.8952 | 8.8952 | 9.0107 | 8.8952 | 8.8952 | 8.8952 | 8.8952
Blank id |  | Initial Reagent Volume | 200 | ml |  | 1 | 2 | H2Om | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
PE/60450 |  | Blank Conc. | 0.6 | ug/ml |  | 20 | 0.4 | LOD | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
Gas Volume |  | = | Vm*(Tr/Tm)*(Pm/Pr) |  |  |  |  | Samp Vol | 309 | 309 | 309 | 309 | 309 | 310.784 | 309 | 309 | 309
|  | = | 222.2407 | dm3 |  |  |  | Samp Con | 240 | 240 | 240 | 240 | 240 | 240 | 253.8564 | 240 | 240
|  | V= | 0.22224 | m3 |  |  |  | Initial Vol | 200 | 200 | 200 | 200 | 200 | 200 | 200 | 201.1547 | 200
Sample Analysis | [7 empty] | Blank Con | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.83094
|  | mass = | concentration * volume
|  | mass in sample = | 74160 | ug |  |  |  | Conc. | 183.1604 | 181.0696 | 184.0535 | 182.1162 | 184.9392 | 184.2196 | 193.7523 | 183.1587 | 183.0461
|  | mass in blank = | 120 | ug | [5 empty] | -2.0908 | 0.89314 | -1.0441 | 1.7789 | 1.0592 | 10.5919 | -0.0017139 | -0.11426
[9 empty] | 122.7473 | 4.3715 | 0.7977 | 1.0902 | 3.1644 | 1.1219 | 112.1885 | 2.9375e-06 | 0.013055
|  | Net mass sampled = | mass in sample - mass in blank
|  | = | 74040 | ug |  |  |  | u | 11.0791
[8 empty] | U | 22.1583
|  | Equiv. mass pollutant sampled = | sampled mass analyte * Mr (p) / Mr (a) / K
|  | M= | 49360 | ug
Concentration
|  | Concentration of pollutant = | M / V
|  | C1 @ ref temperature & pressure = | 222101.5423 | ug/m3
|  | C2 @ ref oxygen = | C1 * (20.9-O2r)/(20.9-O2duct)
|  | = | 183160.3811 | ug/m3
|  | C3 @ ref moisture = | C2 *(100-H2Or)/(100-H2Om)
|  | = | 183160.3811 | ug/m3
Concentration of |  | SO2 = | 183.1604 | ± | 22.1583 | mg/m3
Data Input By: |  | G Thistlethwaite |  | Data Entry Checked By: |  | A Collings
Date: |  | 2001-09-04 |  | Date: |  | 2001-09-09
Reported uncertainty is calculated using a coverage factor of 2, this gives a confidence of approximately 95%

## SOx 1 (2)
Determination of Pollutants |  |  |  | CLIENT | DEFRA |  |  |  | Duct Oxygen, O2m | 5.6953 | %
in Flue Gases by Absorption Method |  |  |  | JOB No. | AEAT/ED49056
| AEAT/ENV/EQ/WI | 100 |  | SITE: | Site 5 |  |  |  |  | Vm | Tm | Pm | O2m | Samp Vol | Samp Con | Initial Vol | Blank Con
Pollutant | SO2 | Mw(p) = | 64 | TEST: | Day 1 |  |  |  | Value | 448.4 | 300 | 100.8 | 5.6953 | 345 | 250 | 200 | 0.6
Analyte | SO4 | Mw(a) = | 96 | DATE: | 2001-07-24 |  |  |  | Abs Error | 8.968 | 2.5 | 1 | 0.2 | 3.45 | 25 | 2 | 0.4
Ratio analyte/pollutant |  | K = | 1 | TEST PERIOD: | 1327-1516 |  |  |  | u | 5.1777 | 1.4434 | 0.57735 | 0.11547 | 1.9919 | 14.4338 | 1.1547 | 0.23094
|  | Ref Temperature, Tr | 0 | oC K | 273
|  | Ref Pressure, Pr | 101.3 | kPa |  |  |  | Mw(p) = | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64
|  | Ref Oxygen, O2r | 11 | % |  | Uncertainty |  | Mw(a) = | 96 | 96 | 96 | 96 | 96 | 96 | 96 | 96 | 96
|  | Ref Moisture, H2Or | 0 | % |  | % | Absolute | K = | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
|  | Meter Reading (start) | 404.3 | dm3
|  | Meter Reading (end) | 852.7 | dm3
|  | Metered Gas Volume, Vm | 448.4 | dm3 |  | 2 | 8.968 | Tr | 273 | 273 | 273 | 273 | 273 | 273 | 273 | 273 | 273
|  | Meter Temperature, Tm | 27 | oC K | 300 |  | 2.5 | Pr | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3
|  | Meter Pressure, Pm | 100.8 | kPa |  |  | 1 | O2r | 11 | 11 | 11 | 11 | 11 | 11 | 11 | 11 | 11
|  | Duct Oxygen, O2duct | 5.6953 | % |  |  | 0.2 | H2Or | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | Meter Moisture, H2Om | 0 | % |  |  |  | Vm | 448.4 | 453.5777 | 448.4 | 448.4 | 448.4 | 448.4 | 448.4 | 448.4 | 448.4
|  | LOD | 0.2 | ug/ml |  |  |  | Tm | 300 | 300 | 301.4434 | 300 | 300 | 300 | 300 | 300 | 300
Sample id |  | Sample Reagent Volume | 345 | ml |  | 1 | 3.45 | Pm | 100.8 | 100.8 | 100.8 | 101.3774 | 100.8 | 100.8 | 100.8 | 100.8 | 100.8
PE/60449 |  | Sample Conc | 250 | ug/ml |  | 10 | 25 | O2m | 5.6953 | 5.6953 | 5.6953 | 5.6953 | 5.8108 | 5.6953 | 5.6953 | 5.6953 | 5.6953
Blank id |  | Initial Reagent Volume | 200 | ml |  | 1 | 2 | H2Om | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
PE/60450 |  | Blank Conc. | 0.6 | ug/ml |  | 20 | 0.4 | LOD | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
Gas Volume |  | = | Vm*(Tr/Tm)*(Pm/Pr) |  |  |  |  | Samp Vol | 345 | 345 | 345 | 345 | 345 | 346.9919 | 345 | 345 | 345
|  | = | 406.03 | dm3 |  |  |  | Samp Con | 250 | 250 | 250 | 250 | 250 | 250 | 264.4338 | 250 | 250
|  | V= | 0.40603 | m3 |  |  |  | Initial Vol | 200 | 200 | 200 | 200 | 200 | 200 | 200 | 201.1547 | 200
Sample Analysis | [7 empty] | Blank Con | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.83094
|  | mass = | concentration * volume
|  | mass in sample = | 86250 | ug |  |  |  | Conc. | 92.0795 | 91.0284 | 92.5225 | 91.5551 | 92.7842 | 92.6119 | 97.4031 | 92.0788 | 92.0301
|  | mass in blank = | 120 | ug | [5 empty] | -1.0511 | 0.44302 | -0.5244 | 0.70464 | 0.53236 | 5.3236 | -0.00074068 | -0.049379
[9 empty] | 30.6994 | 1.1048 | 0.19626 | 0.27499 | 0.49651 | 0.28341 | 28.3409 | 5.486e-07 | 0.0024382
|  | Net mass sampled = | mass in sample - mass in blank
|  | = | 86130 | ug |  |  |  | u | 5.5407
[8 empty] | U | 11.0814
|  | Equiv. mass pollutant sampled = | sampled mass analyte * Mr (p) / Mr (a) / K
|  | M= | 57420 | ug
Concentration
|  | Concentration of pollutant = | M / V
|  | C1 @ ref temperature & pressure = | 141418.1349 | ug/m3
|  | C2 @ ref oxygen = | C1 * (20.9-O2r)/(20.9-O2duct)
|  | = | 92079.5239 | ug/m3
|  | C3 @ ref moisture = | C2 *(100-H2Or)/(100-H2Om)
|  | = | 92079.5239 | ug/m3
Concentration of |  | SO2 = | 92.0795 | ± | 11.0814 | mg/m3
Data Input By: |  | G Thistlethwaite |  | Data Entry Checked By: |  | A Collings
Date: |  | 2001-09-04 |  | Date: |  | 2001-09-09
Reported uncertainty is calculated using a coverage factor of 2, this gives a confidence of approximately 95%

## SOx 2
Determination of Pollutants |  |  |  | CLIENT | DEFRA |  |  |  | Duct Oxygen, O2m | 8.1754 | %
in Flue Gases by Absorption Method |  |  |  | JOB No. | AEAT/ED49056
| AEAT/ENV/EQ/WI | 100 |  | SITE: | Site 5 |  |  |  |  | Vm | Tm | Pm | O2m | Samp Vol | Samp Con | Initial Vol | Blank Con
Pollutant | SO2 | Mw(p) = | 64 | TEST: | Day 2 |  |  |  | Value | 308.6 | 303 | 100.8 | 8.1754 | 351 | 400 | 200 | 0.6
Analyte | SO4 | Mw(a) = | 96 | DATE: | 2001-07-25 |  |  |  | Abs Error | 6.172 | 2.5 | 1 | 0.2 | 3.51 | 40 | 2 | 0.4
Ratio analyte/pollutant |  | K = | 1 | TEST PERIOD: | 1105-1301 |  |  |  | u | 3.5634 | 1.4434 | 0.57735 | 0.11547 | 2.0265 | 23.094 | 1.1547 | 0.23094
|  | Ref Temperature, Tr | 0 | oC K | 273
|  | Ref Pressure, Pr | 101.3 | kPa |  |  |  | Mw(p) = | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64
|  | Ref Oxygen, O2r | 11 | % |  | Uncertainty |  | Mw(a) = | 96 | 96 | 96 | 96 | 96 | 96 | 96 | 96 | 96
|  | Ref Moisture, H2Or | 0 | % |  | % | Absolute | K = | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
|  | Meter Reading (start) | 3853.9 | dm3
|  | Meter Reading (end) | 4162.5 | dm3
|  | Metered Gas Volume, Vm | 308.6 | dm3 |  | 2 | 6.172 | Tr | 273 | 273 | 273 | 273 | 273 | 273 | 273 | 273 | 273
|  | Meter Temperature, Tm | 30 | oC K | 303 |  | 2.5 | Pr | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3
|  | Meter Pressure, Pm | 100.8 | kPa |  |  | 1 | O2r | 11 | 11 | 11 | 11 | 11 | 11 | 11 | 11 | 11
|  | Duct Oxygen, O2duct | 8.1754 | % |  |  | 0.2 | H2Or | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | Meter Moisture, H2Om | 0 | % |  |  |  | Vm | 308.6 | 312.1634 | 308.6 | 308.6 | 308.6 | 308.6 | 308.6 | 308.6 | 308.6
|  | LOD | 0.2 | ug/ml |  |  |  | Tm | 303 | 303 | 304.4434 | 303 | 303 | 303 | 303 | 303 | 303
Sample id |  | Sample Reagent Volume | 351 | ml |  | 1 | 3.51 | Pm | 100.8 | 100.8 | 100.8 | 101.3774 | 100.8 | 100.8 | 100.8 | 100.8 | 100.8
PE/60451 |  | Sample Conc | 400 | ug/ml |  | 10 | 40 | O2m | 8.1754 | 8.1754 | 8.1754 | 8.1754 | 8.2909 | 8.1754 | 8.1754 | 8.1754 | 8.1754
Blank id |  | Initial Reagent Volume | 200 | ml |  | 1 | 2 | H2Om | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
PE/60450 |  | Blank Conc. | 0.6 | ug/ml |  | 20 | 0.4 | LOD | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
Gas Volume |  | = | Vm*(Tr/Tm)*(Pm/Pr) |  |  |  |  | Samp Vol | 351 | 351 | 351 | 351 | 351 | 353.0265 | 351 | 351 | 351
|  | = | 276.6732 | dm3 |  |  |  | Samp Con | 400 | 400 | 400 | 400 | 400 | 400 | 423.094 | 400 | 400
|  | V= | 0.27667 | m3 |  |  |  | Initial Vol | 200 | 200 | 200 | 200 | 200 | 200 | 200 | 201.1547 | 200
Sample Analysis | [7 empty] | Blank Con | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.83094
|  | mass = | concentration * volume
|  | mass in sample = | 140400 | ug |  |  |  | Conc. | 262.984 | 259.982 | 264.2368 | 261.4863 | 265.3923 | 264.5036 | 278.1804 | 262.9827 | 262.8974
|  | mass in blank = | 120 | ug | [5 empty] | -3.002 | 1.2528 | -1.4977 | 2.4083 | 1.5196 | 15.1964 | -0.0012988 | -0.086589
[9 empty] | 251.8713 | 9.0121 | 1.5694 | 2.2431 | 5.8 | 2.3093 | 230.9299 | 1.687e-06 | 0.0074977
|  | Net mass sampled = | mass in sample - mass in blank
|  | = | 140280 | ug |  |  |  | u | 15.8705
[8 empty] | U | 31.7409
|  | Equiv. mass pollutant sampled = | sampled mass analyte * Mr (p) / Mr (a) / K
|  | M= | 93520 | ug
Concentration
|  | Concentration of pollutant = | M / V
|  | C1 @ ref temperature & pressure = | 338016.1658 | ug/m3
|  | C2 @ ref oxygen = | C1 * (20.9-O2r)/(20.9-O2duct)
|  | = | 262984.0115 | ug/m3
|  | C3 @ ref moisture = | C2 *(100-H2Or)/(100-H2Om)
|  | = | 262984.0115 | ug/m3
Concentration of |  | SO2 = | 262.984 | ± | 31.7409 | mg/m3
Data Input By: |  | G Thistlethwaite |  | Data Entry Checked By: |  | A Collings
Date: |  | 2001-09-04 |  | Date: |  | 2001-09-09
Reported uncertainty is calculated using a coverage factor of 2, this gives a confidence of approximately 95%

## SOx 2 (2)
Determination of Pollutants |  |  |  | CLIENT | DEFRA |  |  |  | Duct Oxygen, O2m | 6.6049 | %
in Flue Gases by Absorption Method |  |  |  | JOB No. | AEAT/ED49056
| AEAT/ENV/EQ/WI | 100 |  | SITE: | Site 5 |  |  |  |  | Vm | Tm | Pm | O2m | Samp Vol | Samp Con | Initial Vol | Blank Con
Pollutant | SO2 | Mw(p) = | 64 | TEST: | Day 2 |  |  |  | Value | 306.1 | 304 | 100.8 | 6.6049 | 346 | 52 | 200 | 0.6
Analyte | SO4 | Mw(a) = | 96 | DATE: | 2001-07-25 |  |  |  | Abs Error | 6.122 | 2.5 | 1 | 0.2 | 3.46 | 5.2 | 2 | 0.4
Ratio analyte/pollutant |  | K = | 1 | TEST PERIOD: | 1309-1500 |  |  |  | u | 3.5345 | 1.4434 | 0.57735 | 0.11547 | 1.9976 | 3.0022 | 1.1547 | 0.23094
|  | Ref Temperature, Tr | 0 | oC K | 273
|  | Ref Pressure, Pr | 101.3 | kPa |  |  |  | Mw(p) = | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64
|  | Ref Oxygen, O2r | 11 | % |  | Uncertainty |  | Mw(a) = | 96 | 96 | 96 | 96 | 96 | 96 | 96 | 96 | 96
|  | Ref Moisture, H2Or | 0 | % |  | % | Absolute | K = | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
|  | Meter Reading (start) | 165.9 | dm3
|  | Meter Reading (end) | 472 | dm3
|  | Metered Gas Volume, Vm | 306.1 | dm3 |  | 2 | 6.122 | Tr | 273 | 273 | 273 | 273 | 273 | 273 | 273 | 273 | 273
|  | Meter Temperature, Tm | 31 | oC K | 304 |  | 2.5 | Pr | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3
|  | Meter Pressure, Pm | 100.8 | kPa |  |  | 1 | O2r | 11 | 11 | 11 | 11 | 11 | 11 | 11 | 11 | 11
|  | Duct Oxygen, O2duct | 6.6049 | % |  |  | 0.2 | H2Or | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | Meter Moisture, H2Om | 0 | % |  |  |  | Vm | 306.1 | 309.6345 | 306.1 | 306.1 | 306.1 | 306.1 | 306.1 | 306.1 | 306.1
|  | LOD | 0.2 | ug/ml |  |  |  | Tm | 304 | 304 | 305.4434 | 304 | 304 | 304 | 304 | 304 | 304
Sample id |  | Sample Reagent Volume | 346 | ml |  | 1 | 3.46 | Pm | 100.8 | 100.8 | 100.8 | 101.3774 | 100.8 | 100.8 | 100.8 | 100.8 | 100.8
PE/60452 |  | Sample Conc | 52 | ug/ml |  | 10 | 5.2 | O2m | 6.6049 | 6.6049 | 6.6049 | 6.6049 | 6.7204 | 6.6049 | 6.6049 | 6.6049 | 6.6049
Blank id |  | Initial Reagent Volume | 200 | ml |  | 1 | 2 | H2Om | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
PE/60450 |  | Blank Conc. | 0.6 | ug/ml |  | 20 | 0.4 | LOD | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
Gas Volume |  | = | Vm*(Tr/Tm)*(Pm/Pr) |  |  |  |  | Samp Vol | 346 | 346 | 346 | 346 | 346 | 347.9976 | 346 | 346 | 346
|  | = | 273.5291 | dm3 |  |  |  | Samp Con | 52 | 52 | 52 | 52 | 52 | 52 | 55.0022 | 52 | 52
|  | V= | 0.27353 | m3 |  |  |  | Initial Vol | 200 | 200 | 200 | 200 | 200 | 200 | 200 | 201.1547 | 200
Sample Analysis | [7 empty] | Blank Con | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.83094
|  | mass = | concentration * volume
|  | mass in sample = | 17992 | ug |  |  |  | Conc. | 30.1666 | 29.8223 | 30.3099 | 29.9948 | 30.4123 | 30.342 | 31.92 | 30.1655 | 30.0887
|  | mass in blank = | 120 | ug | [5 empty] | -0.34436 | 0.14323 | -0.1718 | 0.24566 | 0.17534 | 1.7534 | -0.0011694 | -0.077962
[9 empty] | 3.3401 | 0.11858 | 0.020515 | 0.029516 | 0.060348 | 0.030743 | 3.0743 | 1.3676e-06 | 0.0060781
|  | Net mass sampled = | mass in sample - mass in blank
|  | = | 17872 | ug |  |  |  | u | 1.8276
[8 empty] | U | 3.6552
|  | Equiv. mass pollutant sampled = | sampled mass analyte * Mr (p) / Mr (a) / K
|  | M= | 11914.6667 | ug
Concentration
|  | Concentration of pollutant = | M / V
|  | C1 @ ref temperature & pressure = | 43559.0518 | ug/m3
|  | C2 @ ref oxygen = | C1 * (20.9-O2r)/(20.9-O2duct)
|  | = | 30166.6309 | ug/m3
|  | C3 @ ref moisture = | C2 *(100-H2Or)/(100-H2Om)
|  | = | 30166.6309 | ug/m3
Concentration of |  | SO2 = | 30.1666 | ± | 3.6552 | mg/m3
Data Input By: |  | G Thistlethwaite |  | Data Entry Checked By: |  | A Collings
Date: |  | 2001-09-04 |  | Date: |  | 2001-09-09
Reported uncertainty is calculated using a coverage factor of 2, this gives a confidence of approximately 95%

## HCl 1
Determination of Pollutants |  |  |  | CLIENT | DEFRA |  |  |  | Duct Oxygen, O2m | 8.8952 | %
in Flue Gases by Absorption Method |  |  |  | JOB No. | AEAT/ED49056
| AEAT/ENV/EQ/WI | 100 |  | SITE: | Site 5 |  |  |  |  | Vm | Tm | Pm | O2m | Samp Vol | Samp Con | Initial Vol | Blank Con
Pollutant | HCl | Mw(p) = | 36.5 | TEST: | Day 1 |  |  |  | Value | 186.3 | 296 | 100.7 | 8.8952 | 318 | 8.6 | 200 | 0
Analyte | Cl- | Mw(a) = | 35.5 | DATE: | 2001-07-24 |  |  |  | Abs Error | 3.726 | 2.5 | 1 | 0.2 | 3.18 | 0.86 | 2 | 0.01
Ratio analyte/pollutant |  | K = | 1 | TEST PERIOD: | 1125-1318 |  |  |  | u | 2.1512 | 1.4434 | 0.57735 | 0.11547 | 1.836 | 0.49652 | 1.1547 | 0.0057735
|  | Ref Temperature, Tr | 0 | oC K | 273
|  | Ref Pressure, Pr | 101.3 | kPa |  |  |  | Mw(p) = | 36.5 | 36.5 | 36.5 | 36.5 | 36.5 | 36.5 | 36.5 | 36.5 | 36.5
|  | Ref Oxygen, O2r | 11 | % |  | Uncertainty |  | Mw(a) = | 35.5 | 35.5 | 35.5 | 35.5 | 35.5 | 35.5 | 35.5 | 35.5 | 35.5
|  | Ref Moisture, H2Or | 0 | % |  | % | Absolute | K = | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
|  | Meter Reading (start) | 728.4 | dm3
|  | Meter Reading (end) | 914.7 | dm3
|  | Metered Gas Volume, Vm | 186.3 | dm3 |  | 2 | 3.726 | Tr | 273 | 273 | 273 | 273 | 273 | 273 | 273 | 273 | 273
|  | Meter Temperature, Tm | 23 | oC K | 296 |  | 2.5 | Pr | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3
|  | Meter Pressure, Pm | 100.7 | kPa |  |  | 1 | O2r | 11 | 11 | 11 | 11 | 11 | 11 | 11 | 11 | 11
|  | Duct Oxygen, O2duct | 8.8952 | % |  |  | 0.2 | H2Or | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | Meter Moisture, H2Om | 0 | % |  |  |  | Vm | 186.3 | 188.4512 | 186.3 | 186.3 | 186.3 | 186.3 | 186.3 | 186.3 | 186.3
|  | LOD | 0.005 | ug/ml |  |  |  | Tm | 296 | 296 | 297.4434 | 296 | 296 | 296 | 296 | 296 | 296
Sample id |  | Sample Reagent Volume | 318 | ml |  | 1 | 3.18 | Pm | 100.7 | 100.7 | 100.7 | 101.2774 | 100.7 | 100.7 | 100.7 | 100.7 | 100.7
H2O/31139 |  | Sample Conc | 8.6 | ug/ml |  | 10 | 0.86 | O2m | 8.8952 | 8.8952 | 8.8952 | 8.8952 | 9.0107 | 8.8952 | 8.8952 | 8.8952 | 8.8952
Blank id |  | Initial Reagent Volume | 200 | ml |  | 1 | 2 | H2Om | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
H2O/31165 |  | Blank Conc. | 0 | ug/ml |  | 20 | 0.01 | LOD | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005
Gas Volume |  | = | Vm*(Tr/Tm)*(Pm/Pr) |  |  |  |  | Samp Vol | 318 | 318 | 318 | 318 | 318 | 319.836 | 318 | 318 | 318
|  | = | 170.8063 | dm3 |  |  |  | Samp Con | 8.6 | 8.6 | 8.6 | 8.6 | 8.6 | 8.6 | 9.0965 | 8.6 | 8.6
|  | V= | 0.17081 | m3 |  |  |  | Initial Vol | 200 | 200 | 200 | 200 | 200 | 200 | 200 | 201.1547 | 200
Sample Analysis | [7 empty] | Blank Con | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0057735
|  | mass = | concentration * volume
|  | mass in sample = | 2734.8 | ug |  |  |  | Conc. | 13.5758 | 13.4209 | 13.642 | 13.4984 | 13.7077 | 13.6542 | 14.3596 | 13.5758 | 13.5701
|  | mass in blank = | 0 | ug | [5 empty] | -0.15497 | 0.066199 | -0.077392 | 0.13185 | 0.07838 | 0.7838 | 0 | -0.0057321
[9 empty] | 0.67229 | 0.024016 | 0.0043824 | 0.0059894 | 0.017384 | 0.0061434 | 0.61434 | 0 | 3.2856e-05
|  | Net mass sampled = | mass in sample - mass in blank
|  | = | 2734.8 | ug |  |  |  | u | 0.81993
[8 empty] | U | 1.6399
|  | Equiv. mass pollutant sampled = | sampled mass analyte * Mr (p) / Mr (a) / K
|  | M= | 2811.8366 | ug
Concentration
|  | Concentration of pollutant = | M / V
|  | C1 @ ref temperature & pressure = | 16462.1391 | ug/m3
|  | C2 @ ref oxygen = | C1 * (20.9-O2r)/(20.9-O2duct)
|  | = | 13575.825 | ug/m3
|  | C3 @ ref moisture = | C2 *(100-H2Or)/(100-H2Om)
|  | = | 13575.825 | ug/m3
Concentration of |  | HCl = | 13.5758 | ± | 1.6399 | mg/m3
Data Input By: |  | G Thistlethwaite |  | Data Entry Checked By:
Date: |  | 2001-09-04 |  | Date:
Reported uncertainty is calculated using a coverage factor of 2, this gives a confidence of approximately 95%

## HC1 1 (2)
Determination of Pollutants |  |  |  | CLIENT | DEFRA |  |  |  | Duct Oxygen, O2m | 5.6953 | %
in Flue Gases by Absorption Method |  |  |  | JOB No. | AEAT/ED49056
| AEAT/ENV/EQ/WI | 100 |  | SITE: | Site 5 |  |  |  |  | Vm | Tm | Pm | O2m | Samp Vol | Samp Con | Initial Vol | Blank Con
Pollutant | HCl | Mw(p) = | 36.5 | TEST: | Day 1 |  |  |  | Value | 329 | 300 | 100.8 | 5.6953 | 368 | 49 | 200 | 0
Analyte | Cl- | Mw(a) = | 35.5 | DATE: | 2001-07-24 |  |  |  | Abs Error | 6.58 | 2.5 | 1 | 0.2 | 3.68 | 4.9 | 2 | 0.01
Ratio analyte/pollutant |  | K = | 1 | TEST PERIOD: | 1327-1516 |  |  |  | u | 3.799 | 1.4434 | 0.57735 | 0.11547 | 2.1246 | 2.829 | 1.1547 | 0.0057735
|  | Ref Temperature, Tr | 0 | oC K | 273
|  | Ref Pressure, Pr | 101.3 | kPa |  |  |  | Mw(p) = | 36.5 | 36.5 | 36.5 | 36.5 | 36.5 | 36.5 | 36.5 | 36.5 | 36.5
|  | Ref Oxygen, O2r | 11 | % |  | Uncertainty |  | Mw(a) = | 35.5 | 35.5 | 35.5 | 35.5 | 35.5 | 35.5 | 35.5 | 35.5 | 35.5
|  | Ref Moisture, H2Or | 0 | % |  | % | Absolute | K = | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
|  | Meter Reading (start) | 6915.5 | dm3
|  | Meter Reading (end) | 7244.5 | dm3
|  | Metered Gas Volume, Vm | 329 | dm3 |  | 2 | 6.58 | Tr | 273 | 273 | 273 | 273 | 273 | 273 | 273 | 273 | 273
|  | Meter Temperature, Tm | 27 | oC K | 300 |  | 2.5 | Pr | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3
|  | Meter Pressure, Pm | 100.8 | kPa |  |  | 1 | O2r | 11 | 11 | 11 | 11 | 11 | 11 | 11 | 11 | 11
|  | Duct Oxygen, O2duct | 5.6953 | % |  |  | 0.2 | H2Or | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | Meter Moisture, H2Om | 0 | % |  |  |  | Vm | 329 | 332.799 | 329 | 329 | 329 | 329 | 329 | 329 | 329
|  | LOD | 0.005 | ug/ml |  |  |  | Tm | 300 | 300 | 301.4434 | 300 | 300 | 300 | 300 | 300 | 300
Sample id |  | Sample Reagent Volume | 368 | ml |  | 1 | 3.68 | Pm | 100.8 | 100.8 | 100.8 | 101.3774 | 100.8 | 100.8 | 100.8 | 100.8 | 100.8
H2O/31164 |  | Sample Conc | 49 | ug/ml |  | 10 | 4.9 | O2m | 5.6953 | 5.6953 | 5.6953 | 5.6953 | 5.8108 | 5.6953 | 5.6953 | 5.6953 | 5.6953
Blank id |  | Initial Reagent Volume | 200 | ml |  | 1 | 2 | H2Om | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
H2O/31165 |  | Blank Conc. | 0 | ug/ml |  | 20 | 0.01 | LOD | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005
Gas Volume |  | = | Vm*(Tr/Tm)*(Pm/Pr) |  |  |  |  | Samp Vol | 368 | 368 | 368 | 368 | 368 | 370.1246 | 368 | 368 | 368
|  | = | 297.9123 | dm3 |  |  |  | Samp Con | 49 | 49 | 49 | 49 | 49 | 49 | 51.829 | 49 | 49
|  | V= | 0.29791 | m3 |  |  |  | Initial Vol | 200 | 200 | 200 | 200 | 200 | 200 | 200 | 201.1547 | 200
Sample Analysis | [7 empty] | Blank Con | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0057735
|  | mass = | concentration * volume
|  | mass in sample = | 18032 | ug |  |  |  | Conc. | 40.5208 | 40.0582 | 40.7158 | 40.29 | 40.8309 | 40.7547 | 42.8603 | 40.5208 | 40.5182
|  | mass in blank = | 0 | ug | [5 empty] | -0.46255 | 0.19496 | -0.23077 | 0.31009 | 0.23395 | 2.3395 | 0 | -0.0025948
[9 empty] | 5.9292 | 0.21396 | 0.038008 | 0.053254 | 0.096153 | 0.054731 | 5.4731 | 0 | 6.733e-06
|  | Net mass sampled = | mass in sample - mass in blank
|  | = | 18032 | ug |  |  |  | u | 2.435
[8 empty] | U | 4.87
|  | Equiv. mass pollutant sampled = | sampled mass analyte * Mr (p) / Mr (a) / K
|  | M= | 18539.9437 | ug
Concentration
|  | Concentration of pollutant = | M / V
|  | C1 @ ref temperature & pressure = | 62232.8991 | ug/m3
|  | C2 @ ref oxygen = | C1 * (20.9-O2r)/(20.9-O2duct)
|  | = | 40520.7983 | ug/m3
|  | C3 @ ref moisture = | C2 *(100-H2Or)/(100-H2Om)
|  | = | 40520.7983 | ug/m3
Concentration of |  | HCl = | 40.5208 | ± | 4.87 | mg/m3
Data Input By: |  | G Thistlethwaite |  | Data Entry Checked By:
Date: |  | 2001-09-04 |  | Date:
Reported uncertainty is calculated using a coverage factor of 2, this gives a confidence of approximately 95%

## HCl 2
Determination of Pollutants |  |  |  | CLIENT | DEFRA |  |  |  | Duct Oxygen, O2m | 8.1754 | %
in Flue Gases by Absorption Method |  |  |  | JOB No. | AEAT/ED49056
| AEAT/ENV/EQ/WI | 100 |  | SITE: | Site 5 |  |  |  |  | Vm | Tm | Pm | O2m | Samp Vol | Samp Con | Initial Vol | Blank Con
Pollutant | HCl | Mw(p) = | 36.5 | TEST: | Day 2 |  |  |  | Value | 397.1 | 303 | 100.8 | 8.1754 | 333 | 3.1 | 200 | 0
Analyte | Cl- | Mw(a) = | 35.5 | DATE: | 2001-07-25 |  |  |  | Abs Error | 7.942 | 2.5 | 1 | 0.2 | 3.33 | 0.31 | 2 | 0.01
Ratio analyte/pollutant |  | K = | 1 | TEST PERIOD: | 1105-1301 |  |  |  | u | 4.5853 | 1.4434 | 0.57735 | 0.11547 | 1.9226 | 0.17898 | 1.1547 | 0.0057735
|  | Ref Temperature, Tr | 0 | oC K | 273
|  | Ref Pressure, Pr | 101.3 | kPa |  |  |  | Mw(p) = | 36.5 | 36.5 | 36.5 | 36.5 | 36.5 | 36.5 | 36.5 | 36.5 | 36.5
|  | Ref Oxygen, O2r | 11 | % |  | Uncertainty |  | Mw(a) = | 35.5 | 35.5 | 35.5 | 35.5 | 35.5 | 35.5 | 35.5 | 35.5 | 35.5
|  | Ref Moisture, H2Or | 0 | % |  | % | Absolute | K = | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
|  | Meter Reading (start) | 250.7 | dm3
|  | Meter Reading (end) | 647.8 | dm3
|  | Metered Gas Volume, Vm | 397.1 | dm3 |  | 2 | 7.942 | Tr | 273 | 273 | 273 | 273 | 273 | 273 | 273 | 273 | 273
|  | Meter Temperature, Tm | 30 | oC K | 303 |  | 2.5 | Pr | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3
|  | Meter Pressure, Pm | 100.8 | kPa |  |  | 1 | O2r | 11 | 11 | 11 | 11 | 11 | 11 | 11 | 11 | 11
|  | Duct Oxygen, O2duct | 8.1754 | % |  |  | 0.2 | H2Or | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | Meter Moisture, H2Om | 0 | % |  |  |  | Vm | 397.1 | 401.6853 | 397.1 | 397.1 | 397.1 | 397.1 | 397.1 | 397.1 | 397.1
|  | LOD | 0.005 | ug/ml |  |  |  | Tm | 303 | 303 | 304.4434 | 303 | 303 | 303 | 303 | 303 | 303
Sample id |  | Sample Reagent Volume | 333 | ml |  | 1 | 3.33 | Pm | 100.8 | 100.8 | 100.8 | 101.3774 | 100.8 | 100.8 | 100.8 | 100.8 | 100.8
H2O/31166 |  | Sample Conc | 3.1 | ug/ml |  | 10 | 0.31 | O2m | 8.1754 | 8.1754 | 8.1754 | 8.1754 | 8.2909 | 8.1754 | 8.1754 | 8.1754 | 8.1754
Blank id |  | Initial Reagent Volume | 200 | ml |  | 1 | 2 | H2Om | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
H2O/31165 |  | Blank Conc. | 0 | ug/ml |  | 20 | 0.01 | LOD | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005
Gas Volume |  | = | Vm*(Tr/Tm)*(Pm/Pr) |  |  |  |  | Samp Vol | 333 | 333 | 333 | 333 | 333 | 334.9226 | 333 | 333 | 333
|  | = | 356.0172 | dm3 |  |  |  | Samp Con | 3.1 | 3.1 | 3.1 | 3.1 | 3.1 | 3.1 | 3.279 | 3.1 | 3.1
|  | V= | 0.35602 | m3 |  |  |  | Initial Vol | 200 | 200 | 200 | 200 | 200 | 200 | 200 | 201.1547 | 200
Sample Analysis | [7 empty] | Blank Con | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0057735
|  | mass = | concentration * volume
|  | mass in sample = | 1032.3 | ug |  |  |  | Conc. | 2.3195 | 2.293 | 2.3305 | 2.3063 | 2.3407 | 2.3329 | 2.4534 | 2.3195 | 2.3169
|  | mass in blank = | 0 | ug | [5 empty] | -0.026477 | 0.011049 | -0.01321 | 0.021241 | 0.013392 | 0.13392 | 0 | -0.0025945
[9 empty] | 0.019568 | 0.00070105 | 0.00012208 | 0.00017449 | 0.00045118 | 0.00017933 | 0.017933 | 0 | 6.7315e-06
|  | Net mass sampled = | mass in sample - mass in blank
|  | = | 1032.3 | ug |  |  |  | u | 0.13989
[8 empty] | U | 0.27977
|  | Equiv. mass pollutant sampled = | sampled mass analyte * Mr (p) / Mr (a) / K
|  | M= | 1061.3789 | ug
Concentration
|  | Concentration of pollutant = | M / V
|  | C1 @ ref temperature & pressure = | 2981.2572 | ug/m3
|  | C2 @ ref oxygen = | C1 * (20.9-O2r)/(20.9-O2duct)
|  | = | 2319.4837 | ug/m3
|  | C3 @ ref moisture = | C2 *(100-H2Or)/(100-H2Om)
|  | = | 2319.4837 | ug/m3
Concentration of |  | HCl = | 2.3195 | ± | 0.27977 | mg/m3
Data Input By: |  | G Thistlethwaite |  | Data Entry Checked By:
Date: |  | 2001-09-04 |  | Date:
Reported uncertainty is calculated using a coverage factor of 2, this gives a confidence of approximately 95%

## HCl 2 (2)
Determination of Pollutants |  |  |  | CLIENT | DEFRA |  |  |  | Duct Oxygen, O2m | 6.6049 | %
in Flue Gases by Absorption Method |  |  |  | JOB No. | AEAT/ED49056
| AEAT/ENV/EQ/WI | 100 |  | SITE: | Site 5 |  |  |  |  | Vm | Tm | Pm | O2m | Samp Vol | Samp Con | Initial Vol | Blank Con
Pollutant | HCl | Mw(p) = | 36.5 | TEST: | Day 2 |  |  |  | Value | 179.2 | 304 | 100.8 | 6.6049 | 371 | 29 | 200 | 0
Analyte | Cl- | Mw(a) = | 35.5 | DATE: | 2001-07-25 |  |  |  | Abs Error | 3.584 | 2.5 | 1 | 0.2 | 3.71 | 2.9 | 2 | 0.01
Ratio analyte/pollutant |  | K = | 1 | TEST PERIOD: | 1309-1500 |  |  |  | u | 2.0692 | 1.4434 | 0.57735 | 0.11547 | 2.142 | 1.6743 | 1.1547 | 0.0057735
|  | Ref Temperature, Tr | 0 | oC K | 273
|  | Ref Pressure, Pr | 101.3 | kPa |  |  |  | Mw(p) = | 36.5 | 36.5 | 36.5 | 36.5 | 36.5 | 36.5 | 36.5 | 36.5 | 36.5
|  | Ref Oxygen, O2r | 11 | % |  | Uncertainty |  | Mw(a) = | 35.5 | 35.5 | 35.5 | 35.5 | 35.5 | 35.5 | 35.5 | 35.5 | 35.5
|  | Ref Moisture, H2Or | 0 | % |  | % | Absolute | K = | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
|  | Meter Reading (start) | 650.5 | dm3
|  | Meter Reading (end) | 829.7 | dm3
|  | Metered Gas Volume, Vm | 179.2 | dm3 |  | 2 | 3.584 | Tr | 273 | 273 | 273 | 273 | 273 | 273 | 273 | 273 | 273
|  | Meter Temperature, Tm | 31 | oC K | 304 |  | 2.5 | Pr | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3 | 101.3
|  | Meter Pressure, Pm | 100.8 | kPa |  |  | 1 | O2r | 11 | 11 | 11 | 11 | 11 | 11 | 11 | 11 | 11
|  | Duct Oxygen, O2duct | 6.6049 | % |  |  | 0.2 | H2Or | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | Meter Moisture, H2Om | 0 | % |  |  |  | Vm | 179.2 | 181.2692 | 179.2 | 179.2 | 179.2 | 179.2 | 179.2 | 179.2 | 179.2
|  | LOD | 0.005 | ug/ml |  |  |  | Tm | 304 | 304 | 305.4434 | 304 | 304 | 304 | 304 | 304 | 304
Sample id |  | Sample Reagent Volume | 371 | ml |  | 1 | 3.71 | Pm | 100.8 | 100.8 | 100.8 | 101.3774 | 100.8 | 100.8 | 100.8 | 100.8 | 100.8
H2O/31167 |  | Sample Conc | 29 | ug/ml |  | 10 | 2.9 | O2m | 6.6049 | 6.6049 | 6.6049 | 6.6049 | 6.7204 | 6.6049 | 6.6049 | 6.6049 | 6.6049
Blank id |  | Initial Reagent Volume | 200 | ml |  | 1 | 2 | H2Om | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
H2O/31165 |  | Blank Conc. | 0 | ug/ml |  | 20 | 0.01 | LOD | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005 | 0.005
Gas Volume |  | = | Vm*(Tr/Tm)*(Pm/Pr) |  |  |  |  | Samp Vol | 371 | 371 | 371 | 371 | 371 | 373.142 | 371 | 371 | 371
|  | = | 160.132 | dm3 |  |  |  | Samp Con | 29 | 29 | 29 | 29 | 29 | 29 | 30.6743 | 29 | 29
|  | V= | 0.16013 | m3 |  |  |  | Initial Vol | 200 | 200 | 200 | 200 | 200 | 200 | 200 | 201.1547 | 200
Sample Analysis | [7 empty] | Blank Con | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0057735
|  | mass = | concentration * volume
|  | mass in sample = | 10759 | ug |  |  |  | Conc. | 47.8417 | 47.2956 | 48.0689 | 47.5692 | 48.2313 | 48.1179 | 50.6038 | 47.8417 | 47.8366
|  | mass in blank = | 0 | ug | [5 empty] | -0.54612 | 0.22715 | -0.27246 | 0.38959 | 0.27621 | 2.7621 | 0 | -0.0051346
[9 empty] | 8.2816 | 0.29825 | 0.051597 | 0.074235 | 0.15178 | 0.076294 | 7.6294 | 0 | 2.6364e-05
|  | Net mass sampled = | mass in sample - mass in blank
|  | = | 10759 | ug |  |  |  | u | 2.8778
[8 empty] | U | 5.7556
|  | Equiv. mass pollutant sampled = | sampled mass analyte * Mr (p) / Mr (a) / K
|  | M= | 11062.0704 | ug
Concentration
|  | Concentration of pollutant = | M / V
|  | C1 @ ref temperature & pressure = | 69080.944 | ug/m3
|  | C2 @ ref oxygen = | C1 * (20.9-O2r)/(20.9-O2duct)
|  | = | 47841.7058 | ug/m3
|  | C3 @ ref moisture = | C2 *(100-H2Or)/(100-H2Om)
|  | = | 47841.7058 | ug/m3
Concentration of |  | HCl = | 47.8417 | ± | 5.7556 | mg/m3
Data Input By: |  | G Thistlethwaite |  | Data Entry Checked By:
Date: |  | 2001-09-04 |  | Date:
Reported uncertainty is calculated using a coverage factor of 2, this gives a confidence of approximately 95%

## TPM Mst
Determination Of Total Particulate Matter |  |  |  | CLIENT: | DEFRA |  |  |  | Duct Oxygen, O2m | 7.4616
Concentration in Flue Gases |  |  |  | JOB No: | AEAT/ED49056
(Manual sampling Method) |  |  |  | SITE: | Site 5
|  |  |  | TEST: | Day 1
| AEAT/ENV/EQ/WI/ | 132 |  | DATE: | 2001-07-24
|  |  |  | Test Period: | 1119-1522 |  |  |  |  | Vs | O2m | Tm | Pori | M1 | M2 | M3 | M4 | M5 | M6 | M7 | M8
|  | Reference Press, Pr | 101.3 | kPa |  |  |  |  | u | 0.032427 | 0.11547 | 1.4434 | 0.57735 | 0.00017321 | 0.00017321 | 0.00017321 | 0.00017321 | 0.00017321 | 0.00017321 | 0.00017321 | 0.00017321
|  | Reference Temp, Tr | 0 | oC K | 273
|  | Reference Moisture, Xr | 0 | % |  |  |  | Vs | 1.1233 | 1.1557 | 1.1233 | 1.1233 | 1.1233 | 1.1233 | 1.1233 | 1.1233 | 1.1233 | 1.1233 | 1.1233 | 1.1233 | 1.1233
|  | Reference Oxygen, O2r | 11 | % |  | Uncertainty |  | O2r | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | Test Duration, ts | 240 | mins |  | % | Absolute | O2m | 7.4616 | 7.4616 | 7.5771 | 7.4616 | 7.4616 | 7.4616 | 7.4616 | 7.4616 | 7.4616 | 7.4616 | 7.4616 | 7.4616 | 7.4616
|  | Stack Oxygen, O2m | 7.4616 | % |  |  | 0.2 | Xr | 7.8774 | 7.8774 | 7.8774 | 7.8774 | 7.8774 | 7.8774 | 7.8774 | 7.8774 | 7.8774 | 7.8774 | 7.8774 | 7.8774 | 7.8774
|  | Meter Temperature, Tm | 25 | oC | 298 |  | 2.5 | Tm | 298 | 298 | 298 | 299.4434 | 298 | 298 | 298 | 298 | 298 | 298 | 298 | 298 | 298
|  | Average Orfice Pressure, Pori | 2 | mmH2O | 0.01962 |  | 1 | Pori | 0.01962 | 0.01962 | 0.01962 | 0.01962 | 0.59697 | 0.01962 | 0.01962 | 0.01962 | 0.01962 | 0.01962 | 0.01962 | 0.01962 | 0.01962
|  | Barometric Pressure, Pamb | 99.7 | kPa
|  | Volume sampled, Vs | 1.1233 | m3 STP (dry) |  | 5 | 0.056165 | M1 | 0.7675 | 0.7675 | 0.7675 | 0.7675 | 0.7675 | 0.76767 | 0.7675 | 0.7675 | 0.7675 | 0.7675 | 0.7675 | 0.7675 | 0.7675
[8 empty] | M2 | 0.8834 | 0.8834 | 0.8834 | 0.8834 | 0.8834 | 0.8834 | 0.88357 | 0.8834 | 0.8834 | 0.8834 | 0.8834 | 0.8834 | 0.8834
| Sample | Sample | Blank | Blank |  |  |  | M3 | 0.146 | 0.146 | 0.146 | 0.146 | 0.146 | 0.146 | 0.146 | 0.14617 | 0.146 | 0.146 | 0.146 | 0.146 | 0.146
| Filter | Residue | Filter | Residue |  |  |  | M4 | 0.1675 | 0.1675 | 0.1675 | 0.1675 | 0.1675 | 0.1675 | 0.1675 | 0.1675 | 0.16767 | 0.1675 | 0.1675 | 0.1675 | 0.1675
Filter ID: | EPA 400 | ZAM 1822 | EPA 399 | ZAM 1824 |  |  |  | M5 | 0.7686 | 0.7686 | 0.7686 | 0.7686 | 0.7686 | 0.7686 | 0.7686 | 0.7686 | 0.7686 | 0.76877 | 0.7686 | 0.7686 | 0.7686
Filter Weight On: | 0.7675 | 0.146 | 0.7686 | 0.1455 | g | Filter= | 0.0003 | M6 | 0.769 | 0.769 | 0.769 | 0.769 | 0.769 | 0.769 | 0.769 | 0.769 | 0.769 | 0.769 | 0.76917 | 0.769 | 0.769
Filter Weight Off: | 0.8834 | 0.1675 | 0.769 | 0.1452 | g | residue= | 0.0003 | M7 | 0.1455 | 0.1455 | 0.1455 | 0.1455 | 0.1455 | 0.1455 | 0.1455 | 0.1455 | 0.1455 | 0.1455 | 0.1455 | 0.14567 | 0.1455
|  | Weight gain, M | 0.1373 | g |  |  |  | M8 | 0.1452 | 0.1452 | 0.1452 | 0.1452 | 0.1452 | 0.1452 | 0.1452 | 0.1452 | 0.1452 | 0.1452 | 0.1452 | 0.1452 | 0.14537
|  | Duct id | 0.3 | m
|  | Area, Ad | 0.070695 | m2 |  |  |  | M | 0.1373 | 0.1373 | 0.1373 | 0.1373 | 0.1373 | 0.13713 | 0.13747 | 0.13713 | 0.13747 | 0.13747 | 0.13713 | 0.13747 | 0.13713
Gas Volume Sampled
|  | Volume sampled @ STP Vs = | 1.1233 | m3 STP dry
|  | Gas Volume at ref conditions, V = | Vs *(Tr/Tm)*((Pamb+Pori)/Pr)*((20.9-O2m)/(20.9-O2r))
0.91611
|  | (Pm+Pori) = | 99.7196
|  | (Pm+Pori)/Pr = | 0.9844
|  | 20.9-O2m = | 13.4384 |  |  |  |  | V actual | 1.1233 | 1.1233 | 1.1233 | 1.1233 | 1.1233 | 1.1233 | 1.1233 | 1.1233 | 1.1233 | 1.1233 | 1.1233 | 1.1233 | 1.1233
|  | 20.9-O2r = | 9.9 |  |  |  |  | temp corr | 0.91611 | 0.91611 | 0.91611 | 0.91169 | 0.91611 | 0.91611 | 0.91611 | 0.91611 | 0.91611 | 0.91611 | 0.91611 | 0.91611 | 0.91611
|  | O2 Correct = | 1.3574 |  |  |  |  | press corr | 0.9844 | 0.9844 | 0.9844 | 0.9844 | 0.9901 | 0.9844 | 0.9844 | 0.9844 | 0.9844 | 0.9844 | 0.9844 | 0.9844 | 0.9844
|  | V= | 1.3751 | m3 at reference conditions |  |  |  | o2 corr | 0.64298 | 0.64298 | 0.63746 | 0.64298 | 0.64298 | 0.64298 | 0.64298 | 0.64298 | 0.64298 | 0.64298 | 0.64298 | 0.64298 | 0.64298
|  | V= | 1.3751 | m3 STP (ref O2 & H2O) |  |  |  | V | 0.65135 | 0.65135 | 0.64575 | 0.64821 | 0.65512 | 0.65135 | 0.65135 | 0.65135 | 0.65135 | 0.65135 | 0.65135 | 0.65135 | 0.65135
Particle concentration | [7 empty] | Conc | 210.7934 | 210.7934 | 212.6204 | 211.8144 | 209.58 | 210.5275 | 211.0594 | 210.5275 | 211.0594 | 211.0594 | 210.5275 | 211.0594 | 210.5275
|  | Concentration = | M / V | [6 empty] | 0 | 1.827 | 1.021 | -1.2134 | -0.26592 | 0.26592 | -0.26592 | 0.26592 | 0.26592 | -0.26592 | 0.26592 | -0.26592
|  | = | 0.09985 | g.m-3 STP (ref O2 & H2O) |  |  |  | 2.5334 | 6.4182 | 0 | 3.3378 | 1.0424 | 1.4724 | 0.070712 | 0.070712 | 0.070712 | 0.070712 | 0.070712 | 0.070712 | 0.070712 | 0.070712
|  | = | 99.8495 | ± | 9.985 | mg.m-3 STP (ref O2 & H2O)
Data Input By: |  | G Thistlethwaite |  | Data Entry Checked By: |  |  | A Collings
Date: |  | 2001-09-04 |  | Date: |  |  | 2001-09-09
Reported uncertainty is calculated using a coverage factor of 2, this gives a confidence of approximately 95%

## TPM Mst (2)
Determination Of Total Particulate Matter |  |  |  | CLIENT: | DEFRA |  |  |  | Duct Oxygen, O2m | 7.4229
Concentration in Flue Gases |  |  |  | JOB No: | AEAT/ED49056
(Manual sampling Method) |  |  |  | SITE: | Site 5
|  |  |  | TEST: | Day 2
| AEAT/ENV/EQ/WI/ | 132 |  | DATE: | 2001-07-25
|  |  |  | Test Period: | 1101-1500 |  |  |  |  | Vs | O2m | Tm | Pori | M1 | M2 | M3 | M4 | M5 | M6 | M7 | M8
|  | Reference Press, Pr | 101.3 | kPa |  |  |  |  | u | 0.033535 | 0.11547 | 1.4434 | 0.57735 | 0.00017321 | 0.00017321 | 0.00017321 | 0.00017321 | 0.00017321 | 0.00017321 | 0.00017321 | 0.00017321
|  | Reference Temp, Tr | 0 | oC K | 273
|  | Reference Moisture, Xr | 0 | % |  |  |  | Vs | 1.1617 | 1.1952 | 1.1617 | 1.1617 | 1.1617 | 1.1617 | 1.1617 | 1.1617 | 1.1617 | 1.1617 | 1.1617 | 1.1617 | 1.1617
|  | Reference Oxygen, O2r | 11 | % |  | Uncertainty |  | O2r | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | Test Duration, ts | 240 | mins |  | % | Absolute | O2m | 7.4229 | 7.4229 | 7.5384 | 7.4229 | 7.4229 | 7.4229 | 7.4229 | 7.4229 | 7.4229 | 7.4229 | 7.4229 | 7.4229 | 7.4229
|  | Stack Oxygen, O2m | 7.4229 | % |  |  | 0.2 | Xr | 7.8774 | 7.8774 | 7.8774 | 7.8774 | 7.8774 | 7.8774 | 7.8774 | 7.8774 | 7.8774 | 7.8774 | 7.8774 | 7.8774 | 7.8774
|  | Meter Temperature, Tm | 32 | oC | 305 |  | 2.5 | Tm | 305 | 305 | 305 | 306.4434 | 305 | 305 | 305 | 305 | 305 | 305 | 305 | 305 | 305
|  | Average Orfice Pressure, Pori | 2 | mmH2O | 0.01962 |  | 1 | Pori | 0.01962 | 0.01962 | 0.01962 | 0.01962 | 0.59697 | 0.01962 | 0.01962 | 0.01962 | 0.01962 | 0.01962 | 0.01962 | 0.01962 | 0.01962
|  | Barometric Pressure, Pamb | 101.1 | kPa
|  | Volume sampled, Vs | 1.1617 | m3 STP (dry) |  | 5 | 0.058085 | M1 | 0.7703 | 0.7703 | 0.7703 | 0.7703 | 0.7703 | 0.77047 | 0.7703 | 0.7703 | 0.7703 | 0.7703 | 0.7703 | 0.7703 | 0.7703
[8 empty] | M2 | 0.991 | 0.991 | 0.991 | 0.991 | 0.991 | 0.991 | 0.99117 | 0.991 | 0.991 | 0.991 | 0.991 | 0.991 | 0.991
| Sample | Sample | Blank | Blank |  |  |  | M3 | 0.146 | 0.146 | 0.146 | 0.146 | 0.146 | 0.146 | 0.146 | 0.14617 | 0.146 | 0.146 | 0.146 | 0.146 | 0.146
| Filter | Residue | Filter | Residue |  |  |  | M4 | 0.272 | 0.272 | 0.272 | 0.272 | 0.272 | 0.272 | 0.272 | 0.272 | 0.27217 | 0.272 | 0.272 | 0.272 | 0.272
Filter ID: | EPA 401 | ZAM 1823 | EPA 399 | ZAM 1824 |  |  |  | M5 | 0.7686 | 0.7686 | 0.7686 | 0.7686 | 0.7686 | 0.7686 | 0.7686 | 0.7686 | 0.7686 | 0.76877 | 0.7686 | 0.7686 | 0.7686
Filter Weight On: | 0.7703 | 0.146 | 0.7686 | 0.1455 | g | Filter= | 0.0003 | M6 | 0.769 | 0.769 | 0.769 | 0.769 | 0.769 | 0.769 | 0.769 | 0.769 | 0.769 | 0.769 | 0.76917 | 0.769 | 0.769
Filter Weight Off: | 0.991 | 0.272 | 0.769 | 0.1452 | g | residue= | 0.0003 | M7 | 0.1455 | 0.1455 | 0.1455 | 0.1455 | 0.1455 | 0.1455 | 0.1455 | 0.1455 | 0.1455 | 0.1455 | 0.1455 | 0.14567 | 0.1455
|  | Weight gain, M | 0.3466 | g |  |  |  | M8 | 0.1452 | 0.1452 | 0.1452 | 0.1452 | 0.1452 | 0.1452 | 0.1452 | 0.1452 | 0.1452 | 0.1452 | 0.1452 | 0.1452 | 0.14537
|  | Duct id | 0.3 | m
|  | Area, Ad | 0.070695 | m2 |  |  |  | M | 0.3466 | 0.3466 | 0.3466 | 0.3466 | 0.3466 | 0.34643 | 0.34677 | 0.34643 | 0.34677 | 0.34677 | 0.34643 | 0.34677 | 0.34643
Gas Volume Sampled
|  | Volume sampled @ STP Vs = | 1.1617 | m3 STP dry
|  | Gas Volume at ref conditions, V = | Vs *(Tr/Tm)*((Pamb+Pori)/Pr)*((20.9-O2m)/(20.9-O2r))
0.89508
|  | (Pm+Pori) = | 101.1196
|  | (Pm+Pori)/Pr = | 0.99822
|  | 20.9-O2m = | 13.4771 |  |  |  |  | V actual | 1.1617 | 1.1617 | 1.1617 | 1.1617 | 1.1617 | 1.1617 | 1.1617 | 1.1617 | 1.1617 | 1.1617 | 1.1617 | 1.1617 | 1.1617
|  | 20.9-O2r = | 9.9 |  |  |  |  | temp corr | 0.89508 | 0.89508 | 0.89508 | 0.89087 | 0.89508 | 0.89508 | 0.89508 | 0.89508 | 0.89508 | 0.89508 | 0.89508 | 0.89508 | 0.89508
|  | O2 Correct = | 1.3613 |  |  |  |  | press corr | 0.99822 | 0.99822 | 0.99822 | 0.99822 | 1.0039 | 0.99822 | 0.99822 | 0.99822 | 0.99822 | 0.99822 | 0.99822 | 0.99822 | 0.99822
|  | V= | 1.413 | m3 at reference conditions |  |  |  | o2 corr | 0.64484 | 0.64484 | 0.63931 | 0.64484 | 0.64484 | 0.64484 | 0.64484 | 0.64484 | 0.64484 | 0.64484 | 0.64484 | 0.64484 | 0.64484
|  | V= | 1.413 | m3 STP (ref O2 & H2O) |  |  |  | V | 0.66932 | 0.66932 | 0.66358 | 0.66617 | 0.67314 | 0.66932 | 0.66932 | 0.66932 | 0.66932 | 0.66932 | 0.66932 | 0.66932 | 0.66932
Particle concentration | [7 empty] | Conc | 517.8395 | 517.8395 | 522.3147 | 520.2902 | 514.8997 | 517.5808 | 518.0983 | 517.5808 | 518.0983 | 518.0983 | 517.5808 | 518.0983 | 517.5808
|  | Concentration = | M / V | [6 empty] | 0 | 4.4751 | 2.4506 | -2.9399 | -0.25878 | 0.25878 | -0.25878 | 0.25878 | 0.25878 | -0.25878 | 0.25878 | -0.25878
|  | = | 0.24529 | g.m-3 STP (ref O2 & H2O) |  |  |  | 5.9339 | 35.2107 | 0 | 20.0267 | 6.0055 | 8.6428 | 0.066966 | 0.066966 | 0.066966 | 0.066966 | 0.066966 | 0.066966 | 0.066966 | 0.066966
|  | = | 245.2924 | ± | 24.5292 | mg.m-3 STP (ref O2 & H2O)
Data Input By: |  | G Thistlethwaite |  | Data Entry Checked By: |  |  | A Collings
Date: |  | 2001-09-04 |  | Date: |  |  | 2001-09-09
Reported uncertainty is calculated using a coverage factor of 2, this gives a confidence of approximately 95%

## Dioxins Zam
Determination of Dioxin and Furan |  |  |  |  | CLIENT: | DEFRA |  |  |  | Duct Oxygen, O2m | 7.4616
in Flue Gases |  |  |  |  | JOB No: | AEAT/ED49056
Using Zambelli Sampler |  |  |  |  | SITE: | Site 5
[5 empty] | TEST: | Day 1
AEAT/ENV/EQ/WI/105 |  |  |  |  | DATE: | 2001-07-24
[5 empty] | Test Period: | 1119-1523 | [6 empty] | O2m | Vs
|  | Reference Pressure, Pr | = | 101.3 | kPa | [6 empty] | u | 0.11547 | 0.031474
|  | Reference Temp, Tr | = | 0 | deg C K= | 273 | k
|  | Reference Moisture, Xr | = | 0 | % |  |  | Uncertainty |  |  | Xr | 0 | 0 | 0
|  | Reference Oxygen, O2r | = | 11 | % |  |  | % | Absolute |  | O2r | 11 | 11 | 11
|  | Test Duration, ts | = | 240 | mins | [5 empty] | O2m | 7.4616 | 7.5771 | 7.4616
|  | Stack Temp, Td | = | 476 | deg C K= | 749 | k |  | 2.5 |  | Vs | 1.0903 | 1.0903 | 1.1218
|  | Stack Oxygen, O2m | = | 7.4616 | % |  |  |  | 0.2
|  | Static Pressure Duct, Pd | = | 996 | hPa | -0.1 | kPa |  | 1
|  | Barometric Pressure, Pamb | = | 997 | hPa | 99.7 | kPa |  |  |  | V | 1.48 | 1.4673 | 1.5227
|  | Intaken Vol. Counter Litre | = | 1090.3 | Nl (dry) | [7 empty] | -0.012717 | 0.042723
|  | Volume Sampled, Vs | = | 1.0903 | m3 STP dry, stack O2 |  |  | 5 | 0.054515 |  | 0.044576 | 0.001987 | 0.00016172 | 0.0018253
| Filter | STS
Sample ID: | F287 | 3350
Blank ID: |  | 3352
|  | Gas Vol. @ STP dry ref O2 & H2O (V) | = | Vs*(20.9-O2m)/(20.9-O2r)*100/(100-Xr)
|  |  | = | 1.48 | ± | 0.089152 | m3
Congener | LOD | Uncertainty | Mass (ng) |  |  | Uncertainty | ITEQ | Mass | Uncertainty
| ng | % | Sample | Blank | Net | ng | Factor | ng (TEQ) | ng (TEQ)
2378 T4CDD | 0.02 | 100 | 0.069 | 0 | 0.069 | 0.069 | 1 | 0.069 | 0.069
12378 P5CDD | 0.05 | 100 | 0.051 | 0 | 0.051 | 0.051 | 0.5 | 0.0255 | 0.0255
123478 H6CDD | 0.02 | 100 | 0 | 0 | 0.02 | 0 | 0.1 | 0 | 0
123678 H6CDD | 0.02 | 100 | 0 | 0 | 0.02 | 0 | 0.1 | 0 | 0
123789 H6CDD | 0.02 | 100 | 0.041 | 0 | 0.041 | 0.041 | 0.1 | 0.0041 | 0.0041
1234678 H7CDD | 0.04 | 100 | 0.17 | 0.067 | 0.103 | 0.17 | 0.01 | 0.0017 | 0.0017
O8CDD | 0.02 | 25 | 0.51 | 0.27 | 0.24 | 0.1275 | 0.001 | 0.00051 | 0.0001275
2378 T4CDF | 0.01 | 25 | 1 | 0 | 1 | 0.25 | 0.1 | 0.1 | 0.025
12378 P5CDF | 0.009 | 100 | 0.077 | 0 | 0.077 | 0.077 | 0.05 | 0.00385 | 0.00385
23478 P5CDF | 0.009 | 25 | 0.46 | 0.011 | 0.449 | 0.115 | 0.5 | 0.23 | 0.0575 |  | ITEQ
123478 H6CDF | 0.009 | 25 | 0.19 | 0.0094 | 0.1806 | 0.0475 | 0.1 | 0.019 | 0.00475
123678 H6CDF | 0.009 | 25 | 0.15 | 0 | 0.15 | 0.0375 | 0.1 | 0.015 | 0.00375 |  |  |  | M (ng) | V
123789 H6CDF | 0.009 | 100 | 0.021 | 0 | 0.021 | 0.021 | 0.1 | 0.0021 | 0.0021 |  |  | u | 0.056239 | 0.044576
234678 H6CDF | 0.009 | 100 | 0.083 | 0 | 0.083 | 0.083 | 0.1 | 0.0083 | 0.0083
1234678 H7CDF | 0.02 | 25 | 0.46 | 0.024 | 0.436 | 0.115 | 0.01 | 0.0046 | 0.00115 |  | M (ng) | 0.48401 | 0.54025 | 0.48401
1234789 H7CDF | 0.02 | 100 | 0.024 | 0 | 0.024 | 0.024 | 0.01 | 0.00024 | 0.00024 |  | V | 1.48 | 1.48 | 1.5246
O8CDF | 0.03 | 100 | 0.11 | 0.061 | 0.049 | 0.11 | 0.001 | 0.00011 | 0.00011
[11 empty] | C | 0.32704 | 0.36504 | 0.31748
|  |  | Total Dioxins & Dibenzofurans | = | 0.48401 | ± | 0.097408 | ng (TEQ) |  |  |  |  | 0.037999 | -0.0095621
[11 empty] | 0.039184 | 0.0015354 | 0.001444 | 9.1434e-05
|  |  | Concentration of Dioxins & furans | = | 0.32704 | ± | 0.078368 | ng (TEQ).m-3
Mean sampling recovery (%) = |  |  | 75 |  |  | Mean analytical recovery (%)= |  |  | 59
Data Input By: | G Thistlethwaite |  |  |  |  | Data Entry Checked By: |  |  | A Collings
Date: | 2001-09-04 |  |  |  |  | Date: |  |  | 2001-09-09
Reported uncertainty is calculated using a coverage factor of 2, this gives a confidence of approximately 95%
The congeners found in the field blank are provided for information only. They are not subtracted from the sample and are not included in the
calculated total concentration.
Mass of dioxins for the sample, if less than LOD is reported at zero concentration. Field blanks are reported as found.

## Dioxins Zam (max)
Determination of Dioxin and Furan |  |  |  |  | CLIENT: | DEFRA |  |  |  | Duct Oxygen, O2m | 7.4616
in Flue Gases |  |  |  |  | JOB No: | AEAT/ED49056
Using Zambelli Sampler |  |  |  |  | SITE: | Site 5
[5 empty] | TEST: | Day 1
AEAT/ENV/EQ/WI/105 |  |  |  |  | DATE: | 2001-07-24
[5 empty] | Test Period: | 1119-1523 | [6 empty] | O2m | Vs
|  | Reference Pressure, Pr | = | 101.3 | kPa | [6 empty] | u | 0.11547 | 0.031474
|  | Reference Temp, Tr | = | 0 | deg C K= | 273 | k
|  | Reference Moisture, Xr | = | 0 | % |  |  | Uncertainty |  |  | Xr | 0 | 0 | 0
|  | Reference Oxygen, O2r | = | 11 | % |  |  | % | Absolute |  | O2r | 11 | 11 | 11
|  | Test Duration, ts | = | 240 | mins | [5 empty] | O2m | 7.4616 | 7.5771 | 7.4616
|  | Stack Temp, Td | = | 476 | deg C K= | 749 | k |  | 2.5 |  | Vs | 1.0903 | 1.0903 | 1.1218
|  | Stack Oxygen, O2m | = | 7.4616 | % |  |  |  | 0.2
|  | Static Pressure Duct, Pd | = | 996 | hPa | -0.1 | kPa |  | 1
|  | Barometric Pressure, Pamb | = | 997 | hPa | 99.7 | kPa |  |  |  | V | 1.48 | 1.4673 | 1.5227
|  | Intaken Vol. Counter Litre | = | 1090.3 | Nl (dry) | [7 empty] | -0.012717 | 0.042723
|  | Volume Sampled, Vs | = | 1.0903 | m3 STP dry, stack O2 |  |  | 5 | 0.054515 |  | 0.044576 | 0.001987 | 0.00016172 | 0.0018253
| Filter | STS
Sample ID: | F287 | 3350
Blank ID: |  | 3352
|  | Gas Vol. @ STP dry ref O2 & H2O (V) | = | Vs*(20.9-O2m)/(20.9-O2r)*100/(100-Xr)
|  |  | = | 1.48 | ± | 0.089152 | m3
Congener | LOD | Uncertainty | Mass (ng) |  |  | Uncertainty | ITEQ | Mass | Uncertainty
| ng | % | Sample | Blank | Net | ng | Factor | ng (TEQ) | ng (TEQ)
2378 T4CDD | 0.02 | 100 | 0.069 | 0 | 0.069 | 0.069 | 1 | 0.069 | 0.069
12378 P5CDD | 0.05 | 100 | 0.051 | 0 | 0.051 | 0.051 | 0.5 | 0.0255 | 0.0255
123478 H6CDD | 0.02 | 100 | 0.02 | 0 | 0.02 | 0.02 | 0.1 | 0.002 | 0.002
123678 H6CDD | 0.02 | 100 | 0.02 | 0 | 0.02 | 0.02 | 0.1 | 0.002 | 0.002
123789 H6CDD | 0.02 | 100 | 0.041 | 0 | 0.041 | 0.041 | 0.1 | 0.0041 | 0.0041
1234678 H7CDD | 0.04 | 100 | 0.17 | 0.067 | 0.103 | 0.17 | 0.01 | 0.0017 | 0.0017
O8CDD | 0.02 | 25 | 0.51 | 0.27 | 0.24 | 0.1275 | 0.001 | 0.00051 | 0.0001275
2378 T4CDF | 0.01 | 25 | 1 | 0 | 1 | 0.25 | 0.1 | 0.1 | 0.025
12378 P5CDF | 0.009 | 100 | 0.077 | 0 | 0.077 | 0.077 | 0.05 | 0.00385 | 0.00385
23478 P5CDF | 0.009 | 25 | 0.46 | 0.011 | 0.449 | 0.115 | 0.5 | 0.23 | 0.0575 |  | ITEQ
123478 H6CDF | 0.009 | 25 | 0.19 | 0.0094 | 0.1806 | 0.0475 | 0.1 | 0.019 | 0.00475
123678 H6CDF | 0.009 | 25 | 0.15 | 0 | 0.15 | 0.0375 | 0.1 | 0.015 | 0.00375 |  |  |  | M (ng) | V
123789 H6CDF | 0.009 | 100 | 0.021 | 0 | 0.021 | 0.021 | 0.1 | 0.0021 | 0.0021 |  |  | u | 0.056262 | 0.044576
234678 H6CDF | 0.009 | 100 | 0.083 | 0 | 0.083 | 0.083 | 0.1 | 0.0083 | 0.0083
1234678 H7CDF | 0.02 | 25 | 0.46 | 0.024 | 0.436 | 0.115 | 0.01 | 0.0046 | 0.00115 |  | M (ng) | 0.48801 | 0.54427 | 0.48801
1234789 H7CDF | 0.02 | 100 | 0.024 | 0 | 0.024 | 0.024 | 0.01 | 0.00024 | 0.00024 |  | V | 1.48 | 1.48 | 1.5246
O8CDF | 0.03 | 100 | 0.11 | 0.061 | 0.049 | 0.11 | 0.001 | 0.00011 | 0.00011
[11 empty] | C | 0.32974 | 0.36776 | 0.3201
|  |  | Total Dioxins & Dibenzofurans | = | 0.48801 | ± | 0.097449 | ng (TEQ) |  |  |  |  | 0.038015 | -0.0096411
[11 empty] | 0.039219 | 0.0015381 | 0.0014452 | 9.2951e-05
|  |  | Concentration of Dioxins & furans | = | 0.32974 | ± | 0.078438 | ng (TEQ).m-3
Mean sampling recovery (%) = |  |  | 75 |  |  | Mean analytical recovery (%)= |  |  | 59
Data Input By: | G Thistlethwaite |  |  |  |  | Data Entry Checked By: |  |  | A Collings
Date: | 2001-09-04 |  |  |  |  | Date: |  |  | 2001-09-09
Reported uncertainty is calculated using a coverage factor of 2, this gives a confidence of approximately 95%
The congeners found in the field blank are provided for information only. They are not subtracted from the sample and are not included in the
calculated total concentration.
Mass of dioxins for the sample, if less than LOD is reported at LOD concentration. Field blanks are reported as found.

## Dioxins zam 2
Determination of Dioxin and Furan |  |  |  |  | CLIENT: | DEFRA |  |  |  | Duct Oxygen, O2m | 7.4229
in Flue Gases |  |  |  |  | JOB No: | AEAT/ED49056
Using Zambelli Sampler |  |  |  |  | SITE: | Site 5
[5 empty] | TEST: | Day 2
AEAT/ENV/EQ/WI/105 |  |  |  |  | DATE: | 2001-07-25
[5 empty] | Test Period: | 1101-1500 | [6 empty] | O2m | Vs
|  | Reference Pressure, Pr | = | 101.3 | kPa | [6 empty] | u | 0.11547 | 0.031454
|  | Reference Temp, Tr | = | 0 | deg C K= | 273 | k
|  | Reference Moisture, Xr | = | 0 | % |  |  | Uncertainty |  |  | Xr | 0 | 0 | 0
|  | Reference Oxygen, O2r | = | 11 | % |  |  | % | Absolute |  | O2r | 11 | 11 | 11
|  | Test Duration, ts | = | 240 | mins | [5 empty] | O2m | 7.4229 | 7.5384 | 7.4229
|  | Stack Temp, Td | = | 542 | deg C K= | 815 | k |  | 2.5 |  | Vs | 1.0896 | 1.0896 | 1.1211
|  | Stack Oxygen, O2m | = | 7.4229 | % |  |  |  | 0.2
|  | Static Pressure Duct, Pd | = | 1011 | hPa | 0 | kPa |  | 1
|  | Barometric Pressure, Pamb | = | 1011 | hPa | 101.1 | kPa |  |  |  | V | 1.4833 | 1.4706 | 1.5261
|  | Intaken Vol. Counter Litre | = | 1089.6 | Nl (dry) | [7 empty] | -0.012709 | 0.042819
|  | Volume Sampled, Vs | = | 1.0896 | m3 STP dry, stack O2 |  |  | 5 | 0.05448 |  | 0.044665 | 0.001995 | 0.00016151 | 0.0018335
| Filter | STS
Sample ID: | F291 | 3351
Blank ID: | 0 | 3352
|  | Gas Vol. @ STP dry ref O2 & H2O (V) | = | Vs*(20.9-O2m)/(20.9-O2r)*100/(100-Xr)
|  |  | = | 1.4833 | ± | 0.089331 | m3
Congener | LOD | Uncertainty | Mass (ng) |  |  | Uncertainty | ITEQ | Mass | Uncertainty
| ng | % | Sample | Blank | Net | ng | Factor | ng (TEQ) | ng (TEQ)
2378 T4CDD | 0.02 | 100 | 0 | 0 | 0.02 | 0 | 1 | 0 | 0
12378 P5CDD | 0.05 | 100 | 0 | 0 | 0.05 | 0 | 0.5 | 0 | 0
123478 H6CDD | 0.02 | 100 | 0.03 | 0 | 0.03 | 0.03 | 0.1 | 0.003 | 0.003
123678 H6CDD | 0.02 | 100 | 0 | 0 | 0.02 | 0 | 0.1 | 0 | 0
123789 H6CDD | 0.02 | 100 | 0 | 0 | 0.02 | 0 | 0.1 | 0 | 0
1234678 H7CDD | 0.04 | 100 | 0.088 | 0.067 | 0.04 | 0.088 | 0.01 | 0.00088 | 0.00088
O8CDD | 0.02 | 25 | 0.35 | 0.27 | 0.08 | 0.0875 | 0.001 | 0.00035 | 8.75e-05
2378 T4CDF | 0.01 | 25 | 0.27 | 0 | 0.27 | 0.0675 | 0.1 | 0.027 | 0.00675
12378 P5CDF | 0.009 | 100 | 0.066 | 0 | 0.066 | 0.066 | 0.05 | 0.0033 | 0.0033
23478 P5CDF | 0.009 | 25 | 0.15 | 0.011 | 0.139 | 0.0375 | 0.5 | 0.075 | 0.01875 |  | ITEQ
123478 H6CDF | 0.009 | 100 | 0.078 | 0.0094 | 0.0686 | 0.078 | 0.1 | 0.0078 | 0.0078
123678 H6CDF | 0.009 | 100 | 0.071 | 0 | 0.071 | 0.071 | 0.1 | 0.0071 | 0.0071 |  |  |  | M (ng) | V
123789 H6CDF | 0.009 | 100 | 0.012 | 0 | 0.012 | 0.012 | 0.1 | 0.0012 | 0.0012 |  |  | u | 0.013591 | 0.044665
234678 H6CDF | 0.009 | 100 | 0.048 | 0 | 0.048 | 0.048 | 0.1 | 0.0048 | 0.0048
1234678 H7CDF | 0.02 | 25 | 0.27 | 0.024 | 0.246 | 0.0675 | 0.01 | 0.0027 | 0.000675 |  | M (ng) | 0.13406 | 0.14765 | 0.13406
1234789 H7CDF | 0.02 | 100 | 0.043 | 0 | 0.043 | 0.043 | 0.01 | 0.00043 | 0.00043 |  | V | 1.4833 | 1.4833 | 1.528
O8CDF | 0.03 | 25 | 0.5 | 0.061 | 0.439 | 0.125 | 0.001 | 0.0005 | 0.000125
[11 empty] | C | 0.09038 | 0.099543 | 0.087738
|  |  | Total Dioxins & Dibenzofurans | = | 0.13406 | ± | 0.023541 | ng (TEQ) |  |  |  |  | 0.009163 | -0.002642
[11 empty] | 0.0095363 | 9.094e-05 | 8.396e-05 | 6.98e-06
|  |  | Concentration of Dioxins & furans | = | 0.09038 | ± | 0.019073 | ng (TEQ).m-3
Mean sampling recovery (%) = |  |  | 83 |  |  | Mean analytical recovery (%)= |  |  | 58
Data Input By: | G Thistlethwaite |  |  |  |  | Data Entry Checked By: |  |  | A Collings
Date: | 2001-09-04 |  |  |  |  | Date: |  |  | 2001-09-09
Reported uncertainty is calculated using a coverage factor of 2, this gives a confidence of approximately 95%
The congeners found in the field blank are provided for information only. They are not subtracted from the sample and are not included in the
calculated total concentration.
Mass of dioxins for the sample, if less than LOD is reported at zero concentration. Field blanks are reported as found.

## Dioxins zam 2 (max)
Determination of Dioxin and Furan |  |  |  |  | CLIENT: | DEFRA |  |  |  | Duct Oxygen, O2m | 7.4229
in Flue Gases |  |  |  |  | JOB No: | AEAT/ED49056
Using Zambelli Sampler |  |  |  |  | SITE: | Site 5
[5 empty] | TEST: | Day 2
AEAT/ENV/EQ/WI/105 |  |  |  |  | DATE: | 2001-07-25
[5 empty] | Test Period: | 1101-1500 | [6 empty] | O2m | Vs
|  | Reference Pressure, Pr | = | 101.3 | kPa | [6 empty] | u | 0.11547 | 0.031454
|  | Reference Temp, Tr | = | 0 | deg C K= | 273 | k
|  | Reference Moisture, Xr | = | 0 | % |  |  | Uncertainty |  |  | Xr | 0 | 0 | 0
|  | Reference Oxygen, O2r | = | 11 | % |  |  | % | Absolute |  | O2r | 11 | 11 | 11
|  | Test Duration, ts | = | 240 | mins | [5 empty] | O2m | 7.4229 | 7.5384 | 7.4229
|  | Stack Temp, Td | = | 542 | deg C K= | 815 | k |  | 2.5 |  | Vs | 1.0896 | 1.0896 | 1.1211
|  | Stack Oxygen, O2m | = | 7.4229 | % |  |  |  | 0.2
|  | Static Pressure Duct, Pd | = | 1011 | hPa | 0 | kPa |  | 1
|  | Barometric Pressure, Pamb | = | 1011 | hPa | 101.1 | kPa |  |  |  | V | 1.4833 | 1.4706 | 1.5261
|  | Intaken Vol. Counter Litre | = | 1089.6 | Nl (dry) | [7 empty] | -0.012709 | 0.042819
|  | Volume Sampled, Vs | = | 1.0896 | m3 STP dry, stack O2 |  |  | 5 | 0.05448 |  | 0.044665 | 0.001995 | 0.00016151 | 0.0018335
| Filter | STS
Sample ID: | F291 | 3351
Blank ID: | 0 | 3352
|  | Gas Vol. @ STP dry ref O2 & H2O (V) | = | Vs*(20.9-O2m)/(20.9-O2r)*100/(100-Xr)
|  |  | = | 1.4833 | ± | 0.089331 | m3
Congener | LOD | Uncertainty | Mass (ng) |  |  | Uncertainty | ITEQ | Mass | Uncertainty
| ng | % | Sample | Blank | Net | ng | Factor | ng (TEQ) | ng (TEQ)
2378 T4CDD | 0.02 | 100 | 0.02 | 0 | 0.02 | 0.02 | 1 | 0.02 | 0.02
12378 P5CDD | 0.05 | 100 | 0.05 | 0 | 0.05 | 0.05 | 0.5 | 0.025 | 0.025
123478 H6CDD | 0.02 | 100 | 0.03 | 0 | 0.03 | 0.03 | 0.1 | 0.003 | 0.003
123678 H6CDD | 0.02 | 100 | 0.02 | 0 | 0.02 | 0.02 | 0.1 | 0.002 | 0.002
123789 H6CDD | 0.02 | 100 | 0.02 | 0 | 0.02 | 0.02 | 0.1 | 0.002 | 0.002
1234678 H7CDD | 0.04 | 100 | 0.088 | 0.067 | 0.04 | 0.088 | 0.01 | 0.00088 | 0.00088
O8CDD | 0.02 | 25 | 0.35 | 0.27 | 0.08 | 0.0875 | 0.001 | 0.00035 | 8.75e-05
2378 T4CDF | 0.01 | 25 | 0.27 | 0 | 0.27 | 0.0675 | 0.1 | 0.027 | 0.00675
12378 P5CDF | 0.009 | 100 | 0.066 | 0 | 0.066 | 0.066 | 0.05 | 0.0033 | 0.0033
23478 P5CDF | 0.009 | 25 | 0.15 | 0.011 | 0.139 | 0.0375 | 0.5 | 0.075 | 0.01875 |  | ITEQ
123478 H6CDF | 0.009 | 100 | 0.078 | 0.0094 | 0.0686 | 0.078 | 0.1 | 0.0078 | 0.0078
123678 H6CDF | 0.009 | 100 | 0.071 | 0 | 0.071 | 0.071 | 0.1 | 0.0071 | 0.0071 |  |  |  | M (ng) | V
123789 H6CDF | 0.009 | 100 | 0.012 | 0 | 0.012 | 0.012 | 0.1 | 0.0012 | 0.0012 |  |  | u | 0.023001 | 0.044665
234678 H6CDF | 0.009 | 100 | 0.048 | 0 | 0.048 | 0.048 | 0.1 | 0.0048 | 0.0048
1234678 H7CDF | 0.02 | 25 | 0.27 | 0.024 | 0.246 | 0.0675 | 0.01 | 0.0027 | 0.000675 |  | M (ng) | 0.18306 | 0.20606 | 0.18306
1234789 H7CDF | 0.02 | 100 | 0.043 | 0 | 0.043 | 0.043 | 0.01 | 0.00043 | 0.00043 |  | V | 1.4833 | 1.4833 | 1.528
O8CDF | 0.03 | 25 | 0.5 | 0.061 | 0.439 | 0.125 | 0.001 | 0.0005 | 0.000125
[11 empty] | C | 0.12341 | 0.13892 | 0.11981
|  |  | Total Dioxins & Dibenzofurans | = | 0.18306 | ± | 0.039839 | ng (TEQ) |  |  |  |  | 0.015507 | -0.0036076
[11 empty] | 0.015921 | 0.00025348 | 0.00024046 | 1.3015e-05
|  |  | Concentration of Dioxins & furans | = | 0.12341 | ± | 0.031842 | ng (TEQ).m-3
Mean sampling recovery (%) = |  |  | 83 |  |  | Mean analytical recovery (%)= |  |  | 58
Data Input By: | G Thistlethwaite |  |  |  |  | Data Entry Checked By: |  |  | A Collings
Date: | 2001-09-04 |  |  |  |  | Date: |  |  | 2001-09-09
Reported uncertainty is calculated using a coverage factor of 2, this gives a confidence of approximately 95%
The congeners found in the field blank are provided for information only. They are not subtracted from the sample and are not included in the
calculated total concentration.
Mass of dioxins for the sample, if less than LOD is reported at LOD concentration. Field blanks are reported as found.